## CALibrate
This subsystem provides commands for the PSU calibration. Only one channel can be calibrated at a time. If calibration mode has not been enabled with CALibrate:STATe, the calibration commands will generate an error. Use CALibrate:SAVE to save any changes, otherwise all changes will be lost on exit from calibration mode. Within the same calibration session both output voltage and current can be calibrated for the currently selected channel.

SCPI Command	Description
CALibrate[:MODE] {<bool>, <password>}	Enables/disables calibration mode
:CLEar {<password>}	Clears all calibration parameters
:CURRent
[:DATA] {<new value>}	Enters the calibration value
:LEVel {<level>}	Calibrates the output current programming
:PASSword
:NEW {<old>, <new>}	Changes calibration password
:REMark {<string>}	Saves calibration information
:SAVE	Saves the new cal constants in non-volatile memory
:STATe {<bool>, <password>}	Enables calibration parameters
:VOLTage
[:DATA] {<new value>}	Enters the calibration value
:LEVel {<level>}	Calibrates the output voltage programming

### CALibration[:MODE]
Syntax	CALibration[:MODE] {<bool> <password>} / CALibration[:MODE]?
Description	This command enables or disables calibration mode. Calibration mode must be enabled for the channel to accept any calibration commands. The first parameter specifies the ON (1) or OFF (0) state. The second parameter is the password. Successful execution of this command set both output VOLTage and CURRent of the selected channel to the MINimum value (see Section 8.1). / Execution of this command also affects bit 0 (CALibrating) of the Operation Instrument Isummary register (see Section 3.3.2). / If both voltage and current calibration parameters exists on calibration mode exit (CALibration[:MODE] OFF) the CALibration:STATe ON command will automatically follows.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
	<password>	Quoted string	4 to 16 characters	"eezpsu"
Return	The returned parameter is 0 (OFF) or 1 (ON).
Usage example	See Section 10.2
Errors	102,"Invalid cal password" / 104,"Bad sequence of calibration commands"
Related Commands	CALibration:STATe / DIAGnostic[:INFOrmation]:OTIMe?

### CALibration:CLEar
Syntax	CALibration:CLEar {<password>}
Description	Clear all calibration parameters stored in the non-volatile memory for the currently selected channel. After successful execution of this command CALibration:STATe will be set to OFF (0) and further usage of the calibration data will be disabled. This command will be also filled calibration remark with the date and note that calibration data has been cleared.
Parameters	Name	Type	Range	Default
	<password>	Quoted string	4 to 16 characters	"eezpsu"
Return	None
Usage example	DIAG:CAL? / "remark=20161003 new5","u_cal_params_exists=1","u_min_level=0.15V","u_min_data=0.14V","u_min_adc=0.15V","u_mid_level=19.10V","u_mid_data=19.07V","u_mid_adc=18.79V","u_max_level=38.00V","u_max_data=37.90V","u_max_adc=37.39V","u_min_range=0.026046V","u_max_range=39.997372V","i_cal_params_exists=1","i_min_level=0.05A","i_min_data=0.06A","i_min_adc=0.06A","i_mid_level=2.42A","i_mid_data=2.45A","i_mid_adc=2.43A","i_max_level=4.80A","i_max_data=4.84A","i_max_adc=4.81A","i_min_range=0.000000A","i_max_range=5.000022A" / CAL:STAT? / 1 / CAL:CLE “eezpsu” / CAL:STAT? / 0 / DIAG:CAL? / "remark= Not calibrated","u_cal_params_exists=0","i_cal_params_exists=0"
Errors	102,"Invalid cal password"
Related Commands	CALibration:STATe / DIAGnostic[:INFOrmation]:CALibration?

### CALibration:CURRent[:DATA]
Syntax	CALibration:CURRent[:DATA] {<new value>}
Description	This command can only be used when calibration is enabled and the output state of the currently selected channel is ON. It enters a current value that is obtained by reading an external meter. The minimum calibration level (CAL:CURR:LEV MIN) has to be selected first for the value being entered, then the middle and maximum calibration levels (CAL:CURR:LEV MID and CAL:CURR:LEV MAX) for the value being entered. Three successive values must be selected and entered. Data values are expressed in base units – either volts or amperes, depending on which function is being calibrated.
Parameters	Name	Type	Range	Default
	<new value>	numeric (NR2)	-0.2A to MAX + 0.2A / The maximum value is dependent on the PSU current rating. See Section 8.1	–
Return	None
Usage example	See Section 10.2
Errors	104,"Bad sequence of calibration commands" / 107,"Cal value out of range"

### CALibration:CURRent:LEVel
Syntax	CALibration:CURRent:LEVel {<level>}
Description	This command can only be used when calibration is enabled and the output state of the currently selected channel is ON. It sets the PSU to a calibration point that is entered with the CAL:CURR command. During calibration, three points must be entered and the low-end point (MIN) must be selected and entered first. / This command will set output voltage to MAXimum / 2 (for example 25V for the PSU model with 0-50V).
Parameters	Name	Type	Range	Default
	<level>	Discrete	MIN|MID|MAX (see also Section 8.1)	–
Return	None
Usage example	See Section 10.2
Errors	101,"Calibration mode is off" / 104,"Bad sequence of calibration commands"
Related Commands	CALibration:STATe / INSTrument:NSELect / INSTrument[:SELect]

### CALibration:PASSword:NEW
Syntax	CALibration:PASSword:NEW {<old code>, <new code>}
Description	Enter a new calibration password. To change the password, first unsecure the PSU using the old password. Then, the new code has to be entered. The calibration code may contain up to 16 characters over the remote interface. Minimum length is 4 characters. / The new password is automatically stored in nonvolatile memory and does not have to be stored with CALibrate:SAVE.
Parameters	Name	Type	Range	Default
	<old code>	Quoted string	4 to 16 characters	–
	<new code>	Quoted string	4 to 16 characters	–
Return	None
Usage example	See Section 10.2
Errors	102,"Invalid cal password" / 105,"Cal password too long" / 106,"Cal password too short"

### CALibration:REMark
Syntax	CALibration:REMark {<user remark>} / CALibration:REMark?
Description	Record calibration information about the PSU. The calibration message is consist of two parts: / datetime stamp in format yyyymmdd and / up to 32 characters. / The PSU should be in calibration mode before sending a calibration message.
Parameters	Name	Type	Range	Default
	<user remark>	Quoted string	0 to 32 characters	"Calibration passed"
Return	Query the calibration message.
Usage example	When CAL:REM with text "Calibrated by EEZ" is executed at 2015-09-14: / CAL:REM "Calibrated by EEZ" / CAL:REM? / "20150904 Calibrated by EEZ" / See also Section 10.2
Errors	The following errors could be generated by command but not query: / 101,"Calibration mode is off" / 104,"Bad sequence of calibration commands"
Related Commands	CALibration:STATe

### CALibration:SAVE
Syntax	CALibration:SAVE
Description	This command saves calibration constants in non-volatile memory after the calibration procedure has been completed. If calibration mode is exited by programming CALibration:STATe OFF without first saving the new constants, the previous constants are restored. Execution of this command also affects bit 0 (CALibrating) of the Operation Instrument Isummary register (see Section 3.3.2).
Return	None
Usage example	See Section 10.2
Errors	-340,"Calibration failed" / 104,"Bad sequence of calibration commands" / 111,"No new cal data exists"

### CALibration:STATe
Syntax	CALibration:STATe {<bool>} / CALibration:STATe?
Description	This command enables or disables usage of calibration parameters if they exists.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	OFF|ON|0|1	ON
Return	The returned parameter is 0 (OFF) or 1 (ON).
Usage example	CAL:STAT OFF
Errors	110,"Cal params missing or corrupted"
Related Commands	DIAGnostic[:INFOrmation]:OTIMe?

### CALibration:VOLTage[:DATA]
Syntax	CALibration:VOLTage[:DATA] {<new value>}
Description	This command can only be used when calibration is enabled and the output state of the currently selected channel is ON. It enters a voltage value that you obtained by reading an external meter. You must first select the minimum calibration level (CAL:VOLT:LEV MIN) for the value being entered. You must then select the middle and maximum calibration levels (CAL:VOLT:LEV MID and CAL:VOLT:LEV MAX) for the value being entered. Three successive values must be selected and entered. The PSU then computes new voltage calibration constants. These constants has to be stored in non-volatile memory with CALibration:SAVE command.
Parameters	Name	Type	Range	Default
	<new value>	numeric (NR2)	-0.5V to MAX + 0.5V / The maximum value is dependent on the PSU voltage rating. See Section 8.1	–
Return	None
Usage example	See Section Section 10.2
Errors	104,"Bad sequence of calibration commands" / 107,"Cal value out of range"
Related Commands	CALibration:SAVE / CALibration:STATe / INSTrument:NSELect / INSTrument[:SELect]

### CALibration:VOLTage:LEVel
Syntax	CALibration:VOLTage:LEVel {<level>}
Description	This command can only be used when calibration is enabled and the output state of the currently selected channel is ON. It sets the PSU to a calibration point that is entered with the CAL:VOLT[:DATA] command. During calibration, three points must be entered and the low-end point (MIN) must be selected and entered first. / This command will set output current to 50mA.
Parameters	Name	Type	Range	Default
	<level>	Discrete	MIN|MID|MAX (see also Section 8.1)	–
Return	None
Usage example	See Section Section 10.2
Errors	101,"Calibration mode is off" / 104,"Bad sequence of calibration commands"