## INITiate
INITiate commands initialize the trigger system. This enables the trigger system to receive triggers.

SCPI command	Description
INITiate
[:IMMediate]	Completes one full trigger cycle
:CONTinuous {<bool>}	Enables/disables continuous transient triggers

### INITiate
Syntax	INITiate[:IMMediate]
Description	The INITiate command is used to initialize the trigger system of the PSU. This command completes one full trigger cycle when the trigger source is an IMMediate and initiates the trigger subsystem when the trigger source is BUS. / For example, when the IMMediate is selected as a trigger source, an INITiate command / immediately transfers the VOLTage:TRIGgered[:AMPLitude] and CURRent:TRIGgered[:AMPLitude] values to VOLTage[:LEVel][:IMMediate][:AMPLitude] and CURRent[:LEVel][:IMMediate][:AMPLitude] values. Any delay is ignored.
Return	None
Usage example	Generate a trigger operation after 5 seconds: / TRIG:SOUR BUS / TRIG:DEL 5 / INIT / *TRG
Errors	-213,"Init ignored" / 309,"Cannot initiate while in fixed mode"
Related Commands	*TRG / INSTrument:COUPle:TRIGger

### INITiate:CONTinuous
Syntax	INITiate:CONTinuous {<Bool>} / INITiate:CONTinuous?
Description	This command is used to select whether the trigger system is continuously initiated or not. With CONTinuous set to OFF, the trigger system remain in the IDLE state until CONTinuous is set to ON or INITiate:IMMediate is received. / Once CONTinuous is set to ON, the trigger system will be initiated and exit the IDLE state. On completion of each trigger cycle, with CONTinuous ON, the trigger system immediately commence another trigger cycle without entering the IDLE state. / When INITiate:CONTinuous is set to OFF, the current trigger cycle will be completed before entering the IDLE state. The return to IDLE also occur as the result of an ABORt or *RST command. / The ABORt command force the trigger system to the IDLE state; however, the value of INITiate:CONTinuous is unaffected. / If INITiate:CONTinuous was set to ON prior to receiving ABORt, it remains ON and the trigger system immediately exit the IDLE state.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	OFF|ON|0|1	OFF
Return	The query command returns 0 if continuous transients are disabled (OFF), and 1 if continuous transients are enabled (ON).
Usage example	INIT:CONT ON
Related Commands	*RST / ABORt