## SENSe
The SENSe control the current measurement range, energy counting/window, and the data acquisition sequence.

SCPI command	Description
SENSe
:AHOur
:RESet {<channel>}	Resets the amp-hour (Ah) measurement to zero
:CURRent
[:DC]
RANGe[:UPPer] {<range>}	Selects a DC current measurement range
AUTO {<bool>}	Enables/disables seamless measurement auto ranging
:ELOG
:CURRent
[:DC]
:RANge[:UPPer] {<value>}	Selects the Elog current measurement range
:AUTO {<bool>}	Enables/disables Elog seamless measurement autoranging
:FUNCtion
:CURRent {<bool>}	Enables/disables current data logging
:VOLTage {<bool>}	Enables/disables voltage data logging
:PERiod {<time>}	Sets the integration time of an Elog measurement
:WHOur
:RESet {<channel>}	Resets the watt-hour (Wh) measurement to zero

### SENSe:AHOur:RESet
Syntax	SENSe:AHOur:RESet {<channel>}
Description	Use this command to reset the amp-hour counter to zero of the specified channel. If channels are coupled (in series or parallel) or in tracking mode, both channel’s counters will be reset to zero.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SENSE:AHO:RES
Related Commands	FETCh:WHOur? / INSTrument:COUPle:TRACking / OUTPut:TRACk[:STATe] / SENSe:AHOur:RESet / SYSTem:CHANnel:INFOrmation:AHOur:TOTal?

### SENSe:CURRent[:DC]:RANGe[:UPPer]
Syntax	SENSe:CURRent[:DC]:RANGe[:UPPer] {<range>} / SENSe:CURRent[:DC]:RANGe[:UPPer]?
Description	This command selects a DC current measurement range when Power board with multiple current ranges is installed (e.g. r5B12 that can be find out with the SYSTem:CHANnel:MODel? command). The entered value must be higher than the maximum current that you expect to measure. Units are in amperes. / If selected range is lower then measured current, when load is connected, an error -220 will be generated.
Parameters	Name	Type	Range	Default
	<range>	Numeric (NR2), discrete	0.5, 5, MIN|MAX, DEFault	DEF
Return	The query command returns the maximum DC current that can be measured on the range that is presently set.
Usage example	SENS:CURR:RANG? / 0.5
Errors	-220,"Cannot set range"
Related Commands	SENSe:CURRent[:DC]:RANGe:AUTO / SYSTem:CHANnel:MODel?

### SENSe:CURRent[:DC]:RANGe:AUTO
Syntax	SENSe:CURRent[:DC]:RANGe:AUTO {<bool>} / SENSe:CURRent[:DC]:RANGe:AUTO?
Description	This command enables/disables current measurement autoranging. Autoranging provides seamless measurement ranging, which results in a wide dynamic range with no data being lost across ranges. / When enabled, this command overrides the SENSe:CURRent[:DC]:RANGe setting.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
Return	The query command returns 0 if current measurement autoranging is OFF, and 1 if current measurement autoranging is ON.
Usage example	SENS:CURR:RANG:AUTO? / 0
Related Commands	SENSe:CURRent[:DC]:RANGe / SYSTem:CHANnel:MODel?

### SENSe:ELOG[:DC]:RANGe[:UPPer]
Not implemented yet
### SENSe:ELOG[:DC]:RANGe:AUTO
Not implemented yet
### SENSe:ELOG:FUNCtion:CURRent
Not implemented yet
### SENSe:ELOG:FUNCtion:VOLTage
Not implemented yet
### SENSe:ELOG:PERiod
Not implemented yet
### SENSe:WHOur:RESet
Syntax	SENSe:WHOur:RESet {<channel>}
Description	Use this command to reset the watt-hour counter to zero of the specified channel. If channels are coupled (in series or parallel) or in tracking mode, both channel’s counters will be reset to zero.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SENS:WHO:RES
Related Commands	FETCh:AHOur? / INSTrument:COUPle:TRACking / OUTPut:TRACk[:STATe] / SENSe:WHOur:RESet / SYSTem:CHANnel:INFOrmation:WHOur:TOTal