## SYSTem
System commands control system functions that are not directly related to output control, measurement, or status functions.

SCPI command	Description
SYSTem
:BEEPer[:IMMediate]	Issues a single beep immediately
:KEY
:STATe {<bool>}	Enables click tone for local control
:STATe {<bool>}	Enables beeper function
:CAPability?	Returns an <instrument_specifier>
:CHANnel
[:COUNt]?	Returns the number of output channels
:INFOrmation
:CURRent?	Returns output current capability
:ONtime
LAST?	Returns time passed after last output enable
TOTal?	Returns channel’s total active time
:POWer?	Returns output power capability
:PROGram?	Returns programmable features of the channel
:VOLTage?	Returns output voltage capability
:MODel?	Returns the channel model identification
:COMMunicate
:RLSTate {<state>}	Places the instrument in remote or local mode
:CPU
:INFOrmation
:ETHernet
:TYPE?	Returns the type of Ethernet controller
:ONtime
LAST?	Returns time passed after last power on
TOTal?	Returns PSU’s total active time
:TYPE?	Returns the type of CPU
:MODel?	Returns the control board model identification
:OPTion?	Returns information about installed options on the control board
:DATE {<year>,<month>,<day>}	Sets the date of the system clock
:ERRor
[:NEXT]?	Queries and clears errors from the error queue
:COUNt?	Queries the error/event queue for the number of unread items
:KLOCk	Disables front panel [lock/unlock] icon
:LOCal	Places the PSU in the local mode
:PASSword
:CALibration
:RESet	Resets the calibration password to initial value
:FPANel
:RESet	Resets the front panel lock password to initial value
:NEW {<old>, <new>}	Changes system password
:PON:OUTPut:DISable {<bool>}	Sets output state on power up
:POWer {<bool>}	Enters the PSU into the Stand-by mode
:PROTection:TRIP {<bool>}	Enters the PSU into the Stand-by mode in case of protection trip
:REMote	Places the PSU in the remote mode
:RWLock	Places the PSU in the remote mode and disables front panel [lock/unlock] icon
:TEMPerature
:PROTection
[:HIGH]
[:LEVel] {<temperature>[, <sensor>]}	Sets the OTP value
:CLEar[, {<sensor>}]	Clears the latched protection status of the over-temperature protection (OTP)
:DELay
[:TIME] {<delay>[, <sensor>]}	Sets the OTP programming delay
:STATe {<bool>[, <sensor>]}	Enables/disables OTP on the selected temperature sensor
:TRIPped? [<sensor>]	Returns status of OTP activation
:TIME {<hours>,<minutes>,<seconds>}	Sets the time of the system clock
:VERSion?	Returns the SCPI version number

### SYSTem:BEEPer
Syntax	SYSTem:BEEPer[:IMMediate]
Description	This command issues a single beep immediately.
Usage example	SYST:BEEP

### SYSTem:BEEPer:KEY:STATe
Syntax	SYSTem:BEEPer:KEY:STATe {<bool>} (not implemented in M3) / SYSTem:BEEPer:KEY:STATe?
Description	Use this command to enable or disable generation of audiable “click” sound when front panel option is selected.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
Usage example	SYST:BEEP:KEY:STAT ON
Related Commands	SYSTem:BEEPer[:IMMediate] / SYSTem:BEEPer:STATe

### SYSTem:BEEPer:STATe
Syntax	SYSTem:BEEPer:STATe {<bool>} / SYSTem:BEEPer:STATe?
Description	When the beeper is enabled, the PSU generates audible sound in any of the following situations: / the power is turns on or off (see SYSTem:POWer), / when error occurs during front panel operation or remote operation (see Section 7 for the list of error messages), / self-test is failed and / any of the protection function is “tripped”
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	ON
Usage example	SYST:BEEP:STAT ON / SYST:BEEP:STAT? / 1
Related Commands	SYSTem:BEEPer[:IMMediate] / SYSTem:POWer

### SYSTem:CAPability?
Syntax	SYSTem:CAPability?
Description	This query returns the PSU’s capabilities and outputs the appropriate specifiers. / See also SCPI Volume 4: Section 7.1, 1.4.1, 7.2.1, 7.2.2, and 7.2.3
Usage example	SYSTem:CAPability? / DCSUPPLY WITH (MEASURE|MULTIPLE|TRIGGER)

### SYSTem:CHANnel[:COUNt]?
Syntax	SYSTem:CHANnel[:COUNt]?
Description	This query returns the number of output channels in a mainframe.
Usage example	SYSTem:CHANnel? / 2
Related Commands	INSTrument[:SELect] / INSTrument:NSELect

### SYSTem:CHANnel:INFOrmation:CURRent?
Syntax	SYSTem:CHANnel:INFOrmation:CURRent? [<channel>]
Description	Use this query to get currently selected channel output current capability.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:CURR? / 5.00

### SYSTem:CHANnel:INFOrmation:ONTime:LAST?
Syntax	SYSTem:CHANnel:INFOrmation:ONTime:LAST? [<channel>]
Description	This query returns time passed after last activation of the currently selected channel. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:ONT:LAST? / 15m

### SYSTem:CHANnel:INFOrmation:ONTime:TOTal?
Syntax	SYSTem:CHANnel:INFOrmation:ONTime:TOTal? [<channel>]
Description	This query returns total active time of the currently selected channel. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:ONT:TOT? / 1h 45m

### SYSTem:CHANnel:INFOrmation:POWer?
Syntax	SYSTem:CHANnel:INFOrmation:POWEer? [<channel>]
Description	Use this query to get currently selected channel output power capability.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:POW? / 160.00

### SYSTem:CHANnel:INFOrmation:PROGram?
Syntax	SYSTem:CHANnel:INFOrmation:PROGram? [<channel>]
Description	This query returns names of all channel board functionality that can be controlled by firmware. Depending of the board model (see the SYSTem:CHANnel:MODel? query) various combination of the following features can be returned: / Volt – program the output voltage while channel is in the CV mode of operation (see [SOURce[<n>]]:VOLTage and APPLy commands) / Current – program the output current while channel is in the CC mode of operation (see [SOURce[<n>]]:CURRent and APPLy commands) / Power – set max. allowed output power regardless of the channel mode of operation (see [SOURce[<n>]]:POWer:LIMit) / OE – set channel power output (see OUTPut[:STATe]) / Dprog – control down-programmer circuit (see OUTPut:DPRog) / LRipple – set low power mode of operation when SMPS pre-regulator is switched off (see OUTPut:LRIPple) / Rprog – control output voltage programming source (see [SOURce[<n>]]:VOLTage:PROGram[:SOURce])
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:PROG? CH1 / "Volt", "Current", "Power", "OE", "DProg", "LRipple", "Rprog"
Related Commands	APPLy / OUTPut:DPRog / OUTPut:MODE? / OUTPut[:STATe] / OUTPut:LRIPple / [SOURce[<n>]]:CURRent / [SOURce[<n>]]:POWer:LIMit / [SOURce[<n>]]:VOLTage / [SOURce[<n>]]:VOLTage:PROGram[:SOURce]) / SYSTem:CHANnel:MODel?

### SYSTem:CHANnel:INFOrmation:VOLTage?
Syntax	SYSTem:CHANnel:INFOrmation:VOLTage? [<channel>]
Description	Use this query to get currently selected channel output voltage capability.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:VOLT? / 40.00

### SYSTem:CHANnel:MODel?
Syntax	SYSTem:CHANnel:MODel? [<channel>]
Description	This query returns the model identification string of the specified channel.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:MOD? / "Power_r5B6b", "Postreg_r4B43a" / SYST:CHAN:MOD? CH2 / "Postreg_r4B43a"

### SYSTem:COMMunicate:RLSTate
Syntax	SYSTem:COMMunicate:RLSTate {<state>} / SYSTem:COMMunicate:RLSTate?
Description	This command configures the remote/local state of the PSU according to the following settings: / LOCal – The PSU is set to front panel and remote interface control. / REMote – The PSU is set to front panel and remote interface control. / RWLock – The front panel keys are disabled. The PSU can only be controlled via the remote interface. This programmable setting is completely independent from the front panel lock/unlock function that is available from the front panel menu. If you use this command to lock the front panel, the front panel will be unlocked when AC power is cycled. / The LOCal parameter is the same as SYSTem:LOCal, the REMote parameter is the same as SYSTem:REMote, and the RWLock parameter is the same as SYSTem:RWLock. / The remote/local state is unaffected by *RST or any SCPI commands other than SYSTem:COMMunicate:RLState.
Return	The query command returns control state of the PSU that could be LOC, REM, or RWL.
Parameters	Name	Type	Range	Default
	<state>	Discrete	LOCal|REMote|RWLock	LOCal
Usage example	SYST:COMM:RLST? / "LOC"
Related Commands	*RST / SYSTem:LOCal / SYSTem:REMote / SYSTem:RWLock

### SYSTem:CPU:INFOrmation:ETHernet:TYPE?
Syntax	SYSTem:CPU:INFOrmation:ETHernet:TYPE?
Description	This query returns the name of Ethernet controller installed on the control board (Arduino shield). If simulator is used it returns “Simulator”.
Return	The information will be returned as a list of quoted strings.
Usage example	SYST:CPU:INFO:ETH:TYPE? / "W5500"

### SYSTem:CPU:INFOrmation:ONTime:LAST?
Syntax	SYSTem:CPU:INFOrmation:ONTime:LAST? [<channel>]
Description	This query returns time passed after last activation of the PSU. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CPU:INFO:ONT:LAST? / 15m

### SYSTem:CPU:INFOrmation:ONTime:TOTal?
Syntax	SYSTem:CPU:INFOrmation:ONTime:TOTal? [<channel>]
Description	This query returns total active time of the PSU. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CPU:INFO:ONT:TOT? / 1h 45m

### SYSTem:CPU:INFOrmation:TYPE?
Syntax	SYSTem:CPU:INFOrmation:TYPE?
Description	This query returns the name of CPU installed on the control board (Arduino shield). If simulator is used it returns “Simulator”.
Usage example	SYST:CPU:TYPE? / "Due"
Related Commands	*IDN?

### SYSTem:CPU:MODel?
Syntax	SYSTem:CPU:MODel?
Description	This query returns the name of the control board (Arduino shield). If simulator is used it returns “Simulator” and its version.
Usage example	If revision 1 Arduino shield is detected: / SYST:CPU:MODel? / "Arduino", "r1B9" / If firmware is running on simulator: / SYST:CPU:MODel? / "Simulator", "M2.0"
Related Commands	SYSTem:CPU:OPTion?

### SYSTem:CPU:OPTion?
Syntax	SYSTem:CPU:OPTion?
Description	This query returns a list of all installed options on the control board.
Return	The information will be returned as a list of quoted strings.
Usage example	SYST:CPU:MODel? / "BPost", "EEPROM", "RTC", "SDcard", "Ethernet", "USB", "Ext_trig", "Ext_prog", "Watchdog", "Fan"
Related Commands	SYSTem:CPU:MODel?

### SYSTem:DATE
Syntax	SYSTem:DATE {<year>,<month>,<day>} / SYSTem:DATE?
Description	Sets the date of the system clock (RTC). Specify the year, month, and day. / The self-test procedure compare date and time stored in RTC registers with values stored in the non-volatile memory (EEPROM). When the later is greater then former or any of them lost integrity (i.e. any of value is outside allowed range: for example seconds are higher then 60 or months are higher then 12, etc.) self-test will failed. The *TST? will return 1 and detailed report could be queried using the DIAGnostic:TEST? command. / The bit 3 (TIME) of the Questionable Status register will be set (see Section 3.4) if datetime self-test failed or datetime was never set.
Return	Query the current date of the system clock in YYYY, MM, DD format.
Parameters	Name	Type	Range	Default
	<year>	Numeric	2000 – 2099	–
	<month>	Numeric	1 – 12	–
	<day>	Numeric	1 – 31	–
Usage example	SYST:DATE? / 2015, 10, 24
Related Commands	*TST? / DIAGnostic[:INFOrmation]:TEST? / SYSTem:TIME

### SYSTem:ERRor
Syntax	SYSTem:ERRor[:NEXT]?
Description	This query command reads and clear errors from the error queue. A record of up to 20 errors can be stored in the PSU’s error queue. See also “Error Messages” in Section 7. / Errors are retrieved in first-in-first-out (FIFO) order. The first error returned is the first error that was stored. The PSU beeps once each time an error is generated. The error queue is cleared when power has been off or after a *CLS command.
Return	SYSTem:ERRor[:NEXT]? queries and clears the error messages in the error queue. The query returns the number and content of the error message.
Usage example	SYST:ERR? / -113,"Undefined header"
Errors	If more than 20 errors have occurred, the last error stored in the queue (the most recent error) is replaced with: / -350,"Queue overflow" / No additional errors are stored until you remove errors from the queue.
Related Commands	*CLS / *RST / SYSTem:ERRor:COUNt

### SYSTem:ERRor:COUNt?
Syntax	SYSTem:ERRor:COUNt?
Description	This query command queries the error/event queue for the number of unread items. As errors and events may occur at any time, more items may be present in the queue at the time it is actually read.
Usage example	SYST:ERR:COUN? / 10
Related Commands	*CLS / *RST / SYSTem:ERRor[:NEXT]

### SYSTem:KLOCk
Syntax	SYSTem:KLOCk
Description	This command similar to the SYSTem:REMote command disables all front-panel options except for [Lock/Unlock] icon. You can push and hold the [Lock/Unlock] for one second to unlock the front panel. The system password may be needed if it is set. / Lock state is saved in non-volatile memory. Therefore, the front panel remains locked even after AC power is cycled.
Usage example	SYST:KLOC
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:REMote

### SYSTem:LOCal
Syntax	SYSTem:LOCal
Description	This command places the PSU in the local mode during remote operation. All options on the front panel are fully functional.
Usage example	SYST:LOC
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:REMote / SYSTem:RWLock

### SYSTem:PASSword:CALibrate:RESet
Syntax	SYSTem:PASSword:CALibrate:RESet
Description	This command resets the calibration password to the firmware default setting, which is "eezpsu". This command does not reset the system password.
Usage example	SYST:PASS:CAL:RES
Related Commands	CALibration[:MODE] / CALibration:PASSword:NEW

### SYSTem:PASSword:FPANel:RESet
Syntax	SYSTem:PASSword:FPANel:RESet
Description	This command resets the front panel lockout password to the firmware default setting, which is empty space (""). This command does not reset the calibration password.
Usage example	SYST:PASS:FPAN:RES
Related Commands	SYSTem:KLOCk / SYSTem:PASSword:NEW

### SYSTem:PASSword:NEW
Syntax	SYSTem:PASSword:NEW {<old code>, <new>}
Description	Enter a new system password. To change the password, first unsecure the PSU using the old password. Then, the new code has to be entered. The calibration code may contain up to 16 characters over the remote interface. Minimum length is 4 characters. / The new password is automatically stored in non-volatile memory
Parameters	Name	Type	Range	Default
	<old>	Quoted string	0 to 16 characters	–
	<new>	Quoted string	4 to 16 characters	–
Usage example	SYST:PASS:NEW "","mypass2016"
Errors	122,"Invalid sys password" / 125,"Sys password too long" / 126,"Sys password too short"
Related Commands	SYSTem:KLOCk

### SYSTem:PON:OUTPut:DISable
Syntax	SYSTem:PON:OUTPut:DISable {<bool>} / SYSTem:PON:OUTPut:DISable?
Description	This command controls status off all channel outputs on power up. If enabled (ON), all outputs will be disabled regardless of what is stored in user profile selected for auto recall.
Return	Query returns status of forced output disabling on power up.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
Usage example	OUTP? / 1 / SYST:PON:OUTP 1 / (Restart) / OUTP? / 0
Related Commands	MEMory:STATe:RECall:AUTO / SYSTem:POWer

### SYSTem:POWer
Syntax	SYSTem:POWer {<bool>} / SYSTem:POWer?
Description	This command controls powering down and powering up sequence of the AC power inputs. The “Stand-by” indicator (LED_PWR, TLC5925 Out15) will be switched on when the PSU enters the Stand-by mode.
Return	Query returns PSU power Stand-by status.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
Usage example	SYST:POW ON / SYST:POW? / 1
Related Commands	*RST / *TST? / SYSTem:BEEP:STATe

### SYSTem:POWer:PROTection:TRIP
Syntax	SYSTem:POWer:PROTection:TRIP {<bool>} / SYSTem:POWer:PROTection:TRIP?
Description	Use this command to shut down (stand-by) the PSU when any of channel’s protection tripped. The SYSTem:POWer command has to be used to power the PSU on again.
Return	Query returns status of shutdown when tripped mode.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
Usage example	SYST:POW:PROT:TRIP? / 1
Related Commands	SYSTem:POWer

### SYSTem:REMote
Syntax	SYSTem:REMote
Description	Use this command to place the PSU into remote mode for serial (via USB) or Ethernet remote control. All front-panel options are disabled except for [Lock/Unlock] icon. You can push and hold the [Lock/Unlock] for one second to unlock the front panel. The system password may be needed if it is set.
Usage example	SYST:REM
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:LOCal / SYSTem:RWLock

### SYSTem:RWLock
Syntax	SYSTem:RWLock
Description	Places the PSU in the remote mode for serial (via USB) or Ethernet remote control. This command is the same as SYSTem:REMote, except that all front panel options are disabled, including the [Lock/Unlock] icon.
Usage example	SYST:RWL
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:LOCal / SYSTem:REMote

### SYSTem:TEMPerature:PROTection[:HIGH][:LEVel]
Syntax	SYSTem:TEMPerature:PROTection[:HIGH][:LEVel] {<temperature>, <sensor>} / SYSTem:TEMPerature:PROTection[:HIGH][:LEVel]? <sensor>
Description	Set the over-temperature protection (OTP) value in degrees Celsius (oC) of the selected temperature sensor. When the over-temperature protection function of the specified temperature sensor is enabled (SYSTem:TEMPerature:PROTection[:HIGH]:STATe), one of the following action will be performed when the temperature exceeds the over-temperature protection value currently set: / AUX – Switch off power of the main transformer and set bit 4 of the Questionable Status register / CH1, CH2 – Disable channel output (OUTPut OFF) and set bit 4 of the Questionable Instrument Isummary registerIf any of above mentioned temperature sensors cause over-temperature condition an error tone will also follow if beeper is enabled (see SYSTem:BEEPer:STATe). / SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped? command can be used to query whether over-temperature protection occurred on the selected temperature sensor.
Return	Query the over-temperature protection (OTP) value of the selected temperature sensor.
Parameters	Name	Type	Range	Default
	<temperature>	Numeric (NR2)	0 – 100	70
	<sensor>	Discrete	{[AUX]|CH1|CH2}	AUX
Usage example	TEMP:PROT:50.5, AUX / TEMP:PROT? / 50.50
Related Commands	*RST / SYSTem:TEMPerature:PROTection[:HIGH]:STATe / SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped?

### SYSTem:TEMPerature:PROTection[:HIGH]:CLEar
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:CLEar[, {<sensor>}]
Description	This command clears the latched protection status when an over-temperature is detected. / All conditions that generate the fault must be removed before the latched status can be cleared. The output is restored to the state it was in before the fault condition occurred.
Return	None
	Name	Type	Range	Default
	<sensor>	Discrete	{[AUX]CH1|CH2}	AUX
Usage example	TEMP:PROT:CLE
Related Commands	SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped

### SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME]
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME] {<delay>[, <sensor>]} / SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME]?[, {<sensor>}]
Description	This command sets the over-temperature protection delay. The over-temperature protection function will not be triggered during the delay time. After the delay time has expired, the over-temperature protection function will be active. / Programmed values can range from 0 to 300 seconds. See also Section 8.1
Return	The query returns programmed over-temperature protection delay.
Parameters	Name	Type	Range	Default
	<delay>	Numeric	0 – 300 seconds	10
	<sensor>	Discrete	{[AUX]CH1|CH2}	AUX
Usage example	TEMP:PROT:DEL 30, S2
Related Commands	*RST / SYSTem:TEMPerature:PROTection[:HIGH][:LEVel]

### SYSTem:TEMPerature:PROTection[:HIGH]:STATe
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:STATe {<bool>[, <sensor>]} / SYSTem:TEMPerature:PROTection[:HIGH]:STATe?[, {<sensor>}]
Description	This command enables or disables the over-temperature protection (OTP) function. The enabled state is ON (1); the disabled state is OFF (0). If the over-temperature protection function is enabled and the measured output power reach value set by [SOURce[<n>]]:POWer:PROTection[:LEVel] the output is disabled and the Questionable Condition status register OPP bit 10 is set.
Return	The query command returns 0 if the current protection state is OFF, and 1 if the current protection state is ON.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
	<sensor>	Discrete	{[AUX]|CH1|CH2}	AUX
Usage example	TEMP:PROT:STAT? CH1 / 0
Related Commands	*RST / SYSTem:TEMPerature:PROTection[:HIGH]:CLEar

### SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped?
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped?[, {<sensor>}]
Description	Query whether OTP occurred on the selected temperature sensor. When protection is tripped bit 4 (TEMPerature) of the Questionable Status register will be set (see Section 3.4). / The SYSTem:TEMPerature:PROTection[:HIGH]:CLEar command can be send to clear OTP condition caused by the selected temperature sensor.
Return	This query returns a 1 if the protection circuit is tripped and a 0 if it is untripped.
Parameters	Name	Type	Range	Default
	<sensor>	Discrete	{[AUX]|CH1|CH2}	AUX
Usage example	TEMP:PROT:TRIP? / 0
Related Commands	SYSTem:TEMPerature:PROTection[:HIGH]:CLEar / SYSTem:TEMPerature:PROTection[:HIGH]:STATe

### SYSTem:TIME
Syntax	SYSTem:TIME {<hours>,<minutes>,<seconds>} / SYSTem:TIME?
Description	Sets the time of the system clock (RTC). Specify the hours, minutes, and seconds. / The self-test procedure compare date and time stored in RTC registers with values stored in the non-volatile memory (EEPROM). When the later is greater then former or any of them lost integrity (i.e. any of value is outside allowed range: for example seconds are higher then 60 or months are higher then 12, etc.) self-test will failed. The *TST? will return 1 and detailed report could be queried using the DIAGnostic:TEST? command. / The bit 3 (TIME) of the Questionable Status register will be set (see Section 3.4) if datetime self-test failed or datetime was never set.
Return	Query the current time of the system clock in HH, MM, SS format.
Parameters	Name	Type	Range	Default
	<hours>	Numeric	0 – 23	–
	<minutes>	Numeric	0 – 59	–
	<seconds>	Numeric	0 – 59	–
Usage example	SYST:TIME? / 15, 10, 33
Related Commands	*TST? / DIAGnostic[:INFOrmation]:TEST? / SYSTem:DATE

### SYSTem:VERSion?
Syntax	SYSTem:VERSion?
Description	This command returns the version of the SCPI (Standard Commands for Programmable Instruments) standard with which the instrument is in compliance
Return	The command returns a string in the form “YYYY.V”, where YYYY represents the year of the version and V represents a version for that year.
Usage example	SYST:VERS? / 1999.0