## INSTrument
Each channel of the PSU is considered as separate (logical) instrument, which is required by the SCPI standard. The INSTrument subsystem provides a mechanism to identify and select instruments and establish coupling to simplify programming of more channels at once.

SCPI command	Description
INSTrument
[:SELect] {<channel>}	Selects the output to be programmed
:COUPle
:TRACking {<type>}	Selects independent, parallel-tracking, or series-tracking mode
:TRIGger {<mode>}	Selects a coupling between channels trigger systems
:NSELect {<channel_number>}	Selects the output to be programmed

### INSTrument[:SELect]
Syntax	INSTrument[:SELect] {<channel>} / INSTrument[:SELect]?
Description	This command selects the output to be programmed by the output identifier. The outputs of the PSU are considered as two logical instruments. The INSTrument command provides a mechanism to identify and select an output. / When one output is selected, the other output is unavailable for programming until selected. The following commands are affected by the INSTrument command: SOURce, MEASure, and CALibration.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Return	Query returns the currently selected output by the INSTrument[:SELect] or INSTrument:NSELect command. The returned value is CH1 or CH2.
Usage example	INST? / CH1 / INST:SEL? / CH1
Related Commands	INSTrument:NSELect

### INSTrument:COUPle:TRACking
Syntax	INSTrument:COUPle:TRACking {<type>} / INSTrument:COUPle:TRACking?
Description	This command selects how channel’s outputs will be internally connected: / independent (NONE) / parallel-tracking (PARallel) controls K_PAR power relay state when max. output current is doubled (e.g. 10 A instead of 5 A), or / series-tracking (SERies) that controls K_SER power relay state and max. output voltage is doubled (e.g. 80 V instead of 40 V). / When channels are coupled, resulting output will be present on different binding posts as indicated with LED_SP (red) instead of LED_OUT1 (green) and/or LED_OUT2 (green). / Coupled channels will be seen as single channel. Therefore the following commands will affect the both channel regardless of which channel is currently selected: / OUTPut[:STATe], , OUTPut:DPRog, OUTPut:PROTection:CLEar / MEASure[:SCALar]:CURRent[:DC], MEASure[:SCALar]:POWer[:DC], MEASure[:SCALar][:VOLTage][:DC] / [SOURce[<n>]]:CURRent, [SOURce[<n>]]:CURRent:STEP, [SOURce[<n>]]:CURRent:TRIGgered, [SOURce[<n>]]:CURRent:LIMit, [SOURce[<n>]]:CURRent:MODE, [SOURce[<n>]]:CURRent:PROTection:DELay, [SOURce[<n>]]:CURRent:PROTection:STATe, [SOURce[<n>]]:CURRent:PROTection:TRIPped?, [SOURce[<n>]]:LIST:COUNt, [SOURce[<n>]]:LIST:CURRent, [SOURce[<n>]]:LIST:DWELl, [SOURce[<n>]]:LIST:VOLTage[:LEVel], [SOURce[<n>]]:LRIPple, [SOURce[<n>]]:LRIPple:AUTO, [SOURce[<n>]]:POWer:LIMit, [SOURce[<n>]]:POWer:PROTection[:LEVel], [SOURce[<n>]]:POWer:PROTection:DELay[:TIME], [SOURce[<n>]]:POWer:PROTection:STATe, [SOURce[<n>]]:POWer:PROTection:TRIPped?, [SOURce[<n>]]:VOLTage, [SOURce[<n>]]:VOLTage:LIMit, [SOURce[<n>]]:VOLTage:STEP, [SOURce[<n>]]:VOLTage:TRIGgered, [SOURce[<n>]]:VOLTage:MODE, [SOURce[<n>]]:VOLTage:PROTection[:LEVel], [SOURce[<n>]]:VOLTage:PROTection:DELay, [SOURce[<n>]]:VOLTage:PROTection:STATe, [SOURce[<n>]]:VOLTage:PROTection:TRIPped? / TRIGger[:SEQuence][:IMMediate], TRIGger[:SEQuence]:DELay, TRIGger[:SEQuence]:SLOPe, TRIGger[:SEQuence]:SOURce / The following channel specific commands will generate an execution error when channels are coupled: / CALibrate, / DIAGnostic, / [SOURce[<n>]]:VOLTage:PROGram, / [SOURce:]VOLTage:SENSe (series-tracking only) / Channels coupled in parallel could have only one down-programmer circuit active that on the channel 1. State of the channel 2 down-programmer will always be off regardless of what is set with the OUTPut:DPRog command. / If channels are calibrated (CAL:STAT? 1) a max. value of the channel that has lower calibrated value will be used as a reference and multiplied by two. For example, if channel 1 max. voltage is 39.98 and channel 2 max. voltage is 40.00 the new max. value for the SERIes-tracking will become 79.96 V instead of 80.00 V.
Parameters	Name	Type	Range	Default
	<type>	Discrete	NONE|PARAllel|SERies	NONE
Return	Query returns the currently selected output coupling state.
Usage example	INST:COUP:TRAC PARA / VOLT 70.00 / MEAS? / 70.00
Errors	-200,"Execution error"
Related Commands	*SAV / OUTPut:DPRog / OUTPut:TRACk[:STATe]

### INSTrument:COUPle:TRIGger
Syntax	INSTrument:COUPle:TRIGger {<mode>} / INSTrument:COUPle:TRIGger?
Description	This command defines a coupling between channels trigger systems. Use ALL parameter to couple or NONE to remove coupling. / At *RST, trigger systems are uncoupled.
Parameters	Name	Type	Range	Default
	<mode>	Discrete	NONE, ALL	NONE
Return	This query returns the currently coupled output.
Usage example	INST:SEL CH1 / VOLT:TRIG 12 / CURR:TRIG 1.5 / INST:SEL CH2 / VOLT:TRIG 5 / CURR:TRIG MAX / INST:COUP:TRIG ALL / TRIG:SOUR IMM / INIT
Related Commands	*RST / INSTrument:COUPle:TRACking

### INSTrument:NSELect
Syntax	INSTrument:NSELect {<channel_number>} / INSTrument:NSELect?
Description	This command is used in conjunction with the SELect command. It serves the same purpose, except that it uses a numeric value instead of the identifier used in the SELect command.
Parameters	Name	Type	Range	Default
	<channel_number>	Numeric	1|2	–
Return	When queried it returns the logical instrument number of the currently selected PSU channel. Note that the numbering used for logical instruments directly corresponds to the numbers used in status reporting for multiple instruments; specifically the STATus:QUEStionable:INSTrument and STATus:OPERation:INSTrument commands.
Usage example	INST:NSEL 2 / INST:NSEL? / 2
Related Commands	STATus:QUEStionable:INSTrument / STATus:OPERation:INSTrument