## SYSTem
System commands control system functions that are not directly related to output control, measurement, or status functions.

SCPI command	Description
SYSTem
:BEEPer[:IMMediate]	Issues a single beep immediately
:KEY
:STATe {<bool>}	Enables click tone for local control
:STATe {<bool>}	Enables beeper function
:CAPability?	Returns an <instrument_specifier>
:CHANnel
[:COUNt]?	Returns the number of output channels
:INFOrmation
:AHOur
TOTal?	Returns channel’s total delivered energy in Ah
:CURRent?	Returns output current capability
:ONTime
LAST?	Returns time passed after last output enable
TOTal?	Returns channel’s total active time
:POWer?	Returns output power capability
:PROGram?	Returns programmable features of the channel
:VOLTage?	Returns output voltage capability
:WHOur
:TOTal?	Returns channel’s total delivered energy in Wh
:MODel?	Returns the channel model identification
:COMMunicate
:ENABle {<bool>, <interface>}	Enables the remote interface
:ETHernet
:ADDRess {<ip_address>}	Sets the static LAN (IP) address
:CONTrol?	Returns the control connection port
:DHCP {<bool>}	Enables the use of the Dynamic Host Configuration Protocol (DHCP)
:DNS <ip_address>	Sets the IP address of the DNS server.
:GATEway {<ip_address>}	Sets the IP address of the default gateway
:MAC?	Returns the MAC address
:PORT {<number>}	Sets the port number
:SMASk {<mask>}	Sets the static subnet mask
:NTP {<server>}	Set s NTP service server address
:RLSTate {<state>}	Places the instrument in remote or local mode
:SERial
:BAUD {<speed>}	Sets the baud rate (speed)
:PARity {<parity>}	Sets the parity mode
:CPU
:INFOrmation
:ETHernet
:TYPE?	Returns the type of Ethernet controller
:ONtime
LAST?	Returns time passed after last power on
TOTal?	Returns PSU’s total active time
:TYPE?	Returns the type of CPU
:MODel?	Returns the control board model identification
:OPTion?	Returns information about installed options on the control board
:DATE {<year>,<month>,<day>}	Sets the date of the system clock
:ERRor
[:NEXT]?	Queries and clears errors from the error queue
:COUNt?	Queries the error/event queue for the number of unread items
:INHibit?	Queries system inhibit state
:KLOCk	Disables front panel [lock/unlock] icon
:LOCal	Places the PSU in the local mode
:PASSword
:CALibration
:RESet	Resets the calibration password to initial value
:FPANel
:RESet	Resets the front panel lock password to initial value
:NEW {<old>, <new>}	Changes system password
:PON:OUTPut:DISable {<bool>}	Sets output state on power up
:POWer {<bool>}	Enters the PSU into the Stand-by mode
:PROTection:TRIP {<bool>}	Enters the PSU into the Stand-by mode in case of protection trip
:REMote	Places the PSU in the remote mode
:RWLock	Places the PSU in the remote mode and disables front panel [lock/unlock] icon
:TEMPerature
:PROTection
[:HIGH]
[:LEVel] {<temperature>[, <sensor>]}	Sets the OTP value
:CLEar [, {<sensor>}]	Clears the latched protection status of the over-temperature protection (OTP)
:DELay
[:TIME] {<delay>[, <sensor>]}	Sets the OTP programming delay
:STATe {<bool>[, <sensor>]}	Enables/disables OTP on the selected temperature sensor
:TRIPped? [<sensor>]	Returns status of OTP activation
:TIME {<hours>,<minutes>,<seconds>}	Sets the time of the system clock
:DTS {rules}	Defines daylight saving time (DST) rules
:ZONE {zone}	Defines time zone
:VERSion?	Returns the SCPI version number

### SYSTem:BEEPer
Syntax	SYSTem:BEEPer[:IMMediate]
Description	This command issues a single beep immediately.
Usage example	SYST:BEEP

### SYSTem:BEEPer:KEY:STATe
Syntax	SYSTem:BEEPer:KEY:STATe {<bool>} / SYSTem:BEEPer:KEY:STATe?
Description	Use this command to enable or disable generation of audiable “click” sound when front panel option is selected.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
Usage example	SYST:BEEP:KEY:STAT ON
Related Commands	SYSTem:BEEPer[:IMMediate] / SYSTem:BEEPer:STATe

### SYSTem:BEEPer:STATe
Syntax	SYSTem:BEEPer:STATe {<bool>} / SYSTem:BEEPer:STATe?
Description	When the beeper is enabled, the PSU generates audible sound in any of the following situations: / the power is turns on or off (see SYSTem:POWer), / when error occurs during front panel operation or remote operation (see Section 7 for the list of error messages), / self-test is failed and / any of the protection function is “tripped”
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	ON
Usage example	SYST:BEEP:STAT ON / SYST:BEEP:STAT? / 1
Related Commands	SYSTem:BEEPer[:IMMediate] / SYSTem:POWer

### SYSTem:CAPability?
Syntax	SYSTem:CAPability?
Description	This query returns the PSU’s capabilities and outputs the appropriate specifiers. / See also SCPI Volume 4: Section 7.1, 1.4.1, 7.2.1, 7.2.2, and 7.2.3
Usage example	SYSTem:CAPability? / DCSUPPLY WITH (MEASURE|MULTIPLE|TRIGGER)

### SYSTem:CHANnel[:COUNt]?
Syntax	SYSTem:CHANnel[:COUNt]?
Description	This query returns the number of output channels in a mainframe.
Usage example	SYSTem:CHANnel? / 2
Related Commands	INSTrument[:SELect] / INSTrument:NSELect

### SYSTem:CHANnel:INFOrmation:AHOur:TOTal?
Syntax	SYSTem:CHANnel:INFOrmation:AHOur:TOTal? [<channel>]
Description	This query returns total delivered energy in amp-hours (Ah) of the currently selected channel. This information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to the last 10 minutes of calculated energy is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:AHO:TOT? CH1 / 1344.35
Related commands	FETCh:AHOur? / FETCh:WHOur? / SENSe:AHOur:RESet / SYSTem:CHANnel:INFOrmation:WHOur:TOTal?

### SYSTem:CHANnel:INFOrmation:CURRent?
Syntax	SYSTem:CHANnel:INFOrmation:CURRent? [<channel>]
Description	Use this query to get currently selected channel output current capability.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:CURR? CH2 / 5.00

### SYSTem:CHANnel:INFOrmation:ONTime:LAST?
Syntax	SYSTem:CHANnel:INFOrmation:ONTime:LAST? [<channel>]
Description	This query returns time passed after last activation of the currently selected channel. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:ONT:LAST? / 15m

### SYSTem:CHANnel:INFOrmation:ONTime:TOTal?
Syntax	SYSTem:CHANnel:INFOrmation:ONTime:TOTal? [<channel>]
Description	This query returns total active time of the currently selected channel. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:ONT:TOT? / 1h 45m

### SYSTem:CHANnel:INFOrmation:POWer?
Syntax	SYSTem:CHANnel:INFOrmation:POWer? [<channel>]
Description	Use this query to get currently selected channel output power capability.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:POW? / 160.00

### SYSTem:CHANnel:INFOrmation:PROGram?
Syntax	SYSTem:CHANnel:INFOrmation:PROGram? [<channel>]
Description	This query returns names of all channel board functionality that can be controlled by firmware. Depending of the board model (see the SYSTem:CHANnel:MODel? query) various combination of the following features can be returned: / Volt – program the output voltage while channel is in the CV mode of operation (see [SOURce[<n>]]:VOLTage and APPLy commands) / Current – program the output current while channel is in the CC mode of operation (see [SOURce[<n>]]:CURRent and APPLy commands) / Power – set max. allowed output power regardless of the channel mode of operation (see [SOURce[<n>]]:POWer:LIMit) / OE – set channel power output (see OUTPut[:STATe]) / Dprog – control down-programmer circuit (see OUTPut:DPRog) / LRipple – set low power mode of operation when SMPS pre-regulator is switched off (see OUTPut:LRIPple) / Rprog – control output voltage programming source (see [SOURce[<n>]]:VOLTage:PROGram[:SOURce])
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:PROG? CH1 / "Volt", "Current", "Power", "OE", "DProg", "LRipple", "Rprog"
Related Commands	APPLy / OUTPut:DPRog / OUTPut:MODE? / OUTPut[:STATe] / OUTPut:LRIPple / [SOURce[<n>]]:CURRent / [SOURce[<n>]]:POWer:LIMit / [SOURce[<n>]]:VOLTage / [SOURce[<n>]]:VOLTage:PROGram[:SOURce]) / SYSTem:CHANnel:MODel?

### SYSTem:CHANnel:INFOrmation:VOLTage?
Syntax	SYSTem:CHANnel:INFOrmation:VOLTage? [<channel>]
Description	Use this query to get currently selected channel output voltage capability.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:VOLT? / 40.00

### SYSTem:CHANnel:INFOrmation:WHOur:TOTal?
Syntax	SYSTem:CHANnel:INFOrmation:AHOur:TOTal? [<channel>]
Description	This query returns total delivered energy in watt-hours (Wh) of the currently selected channel. This information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to the last 10 minutes of calculated energy is lost after restart caused with power outage or system reset. / Energy specified in Watt hours can be easily converted into e.g. Joules (1 Wh = 3600 J).
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:INFO:WHO:TOT? CH1 / 2044.15
Related commands	FETCh:AHOur? / FETCh:WHOur? / SENSe:WHOur:RESet / SYSTem:CHANnel:INFOrmation:AHOur:TOTal?

### SYSTem:CHANnel:MODel?
Syntax	SYSTem:CHANnel:MODel? [<channel>]
Description	This query returns the model identification string of the specified channel.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CHAN:MOD? / "Power_r5B12"
Related Commands	SYSTem:CPU:MODel?

### SYSTem:COMMunicate:ENABle
Syntax	SYSTem:COMMunicate:ENABle {<bool>, <interface>} / SYSTem:COMMunicate:ENABle? {<interface>}
Description	Enables or disables the remote interface Serial (via USB), Ethernet, NTP service, or the remote service Sockets. The setting is effective after rebooting the PSU. This command setting is not changed by power off or the *RST command. / Remote service Sockets is not supported yet
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	–
	<interface>	Discrete	SERial|ETHernet|NTP|SOCKets	–
Return	This query returns the status of the selected interface that could be 0 (OFF) or 1 (ON).
Usage example	SYST:COMM:ENAB 1, ETH / SYST:COMM:ENAB 0, SER

### SYSTem:COMMunicate:ETHernet:ADDRess
Syntax	SYSTem:COMMunicate:ETHernet:ADDRess {<ip_address>} / SYSTem:COMMunicate:ETHernet:ADDRess?
Description	Set the IP address of the system manually if the DHCP mode is not enabled. If the DHCP mode is enabled setting the IP address using this command will be ignored.
Parameters	Name	Type	Range	Default
	<ip_address>	Quoted string	15 characters max. formatted as four groups of up to 3 digits (range 0–255), separated by “.”	–
Return	The query returns the current IP address sets manually or assigned by the DHCP server.
Usage example	SYST:COMM:ETH:ADDR? / "192.168.10.100"
Related Commands	SYSTem:COMMunicate:ETHernet:DHCP

### SYSTem:COMMunicate:ETHernet:DHCP
Syntax	SYSTem:COMMunicate:ETHernet:DHCP {<bool>} / SYSTem:COMMunicate:ETHernet:DHCP?
Description	Enable or disable the DHCP mode. In DHCP mode, the DHCP server in the current network assigns network parameters (IP ADdress, DNS address, GATEway address and the Subnet MASk) for the PSU.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	ON
Return	Query the status of the DHCP mode.
Usage example	SYST:COMM:ETH:DHCP? / 1

### SYSTem:COMMunicate:ETHernet:DNS
Syntax	SYSTem:COMMunicate:ETHernet:DNS {<ip_address>} / SYSTem:COMMunicate:ETHernet:DNS?
Description	Set the DNS (Domain Name Service) address if the DHCP mode is not enabled. If the DHCP mode is enabled setting the DNS address using this command will be ignored.
Parameters	Name	Type	Range	Default
	<ip_address>	Quoted string	15 characters max. formatted as four groups of up to 3 digits (range 0–255), separated by “.”	–
Return	The query returns the DNS address sets manually or assigned by the DHCP server. / If DHCP is used query will return “unknown”.
Usage example	SYST:COMM:ETH:DNS "192.168.1.200"
Related Commands	SYSTem:COMMunicate:ETHernet:DHCP

### SYSTem:COMMunicate:ETHernet:GATEway
Syntax	SYSTem:COMMunicate:ETHernet:GATEway {<ip_address>} / SYSTem:COMMunicate:ETHernet:GATEway?
Description	Set the network gateway address if the DHCP mode is not enabled. If the DHCP mode is enabled setting the network gateway address using this command will be ignored.
Parameters	Name	Type	Range	Default
	<ip_address>	Quoted string	15 characters max. formatted as four groups of up to 3 digits (range 0–255), separated by “.”	–
Return	The query returns the network gateway address sets manually or assigned by the DHCP server. If DHCP is used query will return “unknown”.
Usage example	SYST:COMM:ETH:GATE? / "192.168.10.1"
Related Commands	SYSTem:COMMunicate:ETHernet:DHCP

### SYSTem:COMMunicate:ETHernet:MAC
Syntax	SYSTem:COMMunicate:ETHernet:MAC {<mac_address>} / SYSTem:COMMunicate:ETHernet:MAC?
Description	Use this command to set Ethernet controller IC (W5500) MAC. You can set any combination of six hexadecimal values separated by “-”. Ethernet will work as long as two different machine in the LAN don't have the same MAC address.
Parameters	Name	Type	Range	Default
	<mac_address>	Quoted string	17 characters formatted as six groups of 2 digits separated by “-”	74-69-69-2D-30-00
Return	The query returns the MAC address as a quoted string (six hexadecimal values separated by “-”).
Usage example	SYST:COMM:ETH:MAC? / "70-60-50-40-30-20"

### SYSTem:COMMunicate:ETHernet:PORT
Syntax	SYSTem:COMMunicate:ETHernet:PORT {<number>} / SYSTem:COMMunicate:ETHernet:PORT?
Description	Use this command to change default (5025) Ethernet communication port for SCPI.
Parameters	Name	Type	Range	Default
	<number>	Numeric (NR1)	1-65535	5025
Return	The query returns the Ethernet communication port number.
Usage example	SYST:COMM:ETH:PORT? / 5025
Related Commands	SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:ETHernet:CONTrol?

### SYSTem:COMMunicate:ETHernet:SMASk
Syntax	SYSTem:COMMunicate:ETHernet:SMASk {<mask>} / SYSTem:COMMunicate:ETHernet:SMASk?
Description	Set the subnet mask if the DHCP mode is not enabled. If the DHCP mode is enabled setting the network gateway address using this command will be ignored.
Parameters	Name	Type	Range	Default
	<mask>	Quoted string	15 characters max. formatted as four groups of up to 3 digits (range 0–255), separated by “.”	–
Return	The query returns the subnet mask sets manually or assigned by the DHCP server. If DHCP is used query will return “unknown”.
Usage example	SYST:COMM:ETH:SMAS "255.255.255.0"
Related Commands	SYSTem:COMMunicate:ETHernet:DHCP

### SYSTem:COMMunicate:NTP
Syntax	SYSTem:COMMunicate:NTP {<server>} / SYSTem:COMMunicate:NTP?
Description	Use this command to set the NTP service server network address. The PSU will try to establish connection with selected NTP service on each power up (hard reset), when *RST is issued or once per day (24h as defined with CONF_NTP_PERIOD_SEC parameter in firmware)
Parameters	Name	Type	Range	Default
	<server>	Quoted string	IP address or NTP server host name up to 32 characters	europe.pool.ntp.org
Return	The query command returns the NTP service server network address.
Usage example	SYST:COMM:ENAB NTP / SYST:COMM:NTP "ntp.ubuntu.com"
Related Commands	*RST / SYSTem:COMMunicate:ENABle / SYSTem:DATE / SYSTem:TIME

### SYSTem:COMMunicate:RLSTate
Syntax	SYSTem:COMMunicate:RLSTate {<state>} / SYSTem:COMMunicate:RLSTate?
Description	This command configures the remote/local state of the PSU according to the following settings: / LOCal – The PSU is set to front panel and remote interface control. / REMote – The PSU is set to front panel and remote interface control. / RWLock – The front panel keys are disabled. The PSU can only be controlled via the remote interface. This programmable setting is completely independent from the front panel lock/unlock function that is available from the front panel menu. If you use this command to lock the front panel, the front panel will be unlocked when AC power is cycled. / The LOCal parameter is the same as SYSTem:LOCal, the REMote parameter is the same as SYSTem:REMote, and the RWLock parameter is the same as SYSTem:RWLock. / The remote/local state is unaffected by *RST or any SCPI commands other than SYSTem:COMMunicate:RLState.
Parameters	Name	Type	Range	Default
	<state>	Discrete	LOCal|REMote|RWLock	LOCal
Return	The query command returns control state of the PSU that could be LOC, REM, or RWL.
Usage example	SYST:COMM:RLST? / "LOC"
Related Commands	*RST / SYSTem:LOCal / SYSTem:REMote / SYSTem:RWLock

### SYSTem:COMMunicate:SERial:BAUD
Syntax	SYSTem:COMMunicate:SERial:BAUD {<speed>} / SYSTem:COMMunicate:SERial:BAUD?
Description	Set the baud rate of the serial (via USB) interface and the unit is Baud.
Parameters	Name	Type	Range	Default
	<speed>	Numeric (NR1)	4800|9600|19200|38400|57600|115200	9600
Return	Query the baud rate of the serial interface.
Usage example	SYST:COMM:SER:BAUD 38400 / SYST:COMM:SER:BAUD? / 38400
Related Commands	SYSTem:COMMunicate:ENABle

### SYSTem:COMMunicate:SERial:PARity
Syntax	SYSTem:COMMunicate:SERial:PARity {<parity>} / SYSTem:COMMunicate:SERial:PARity?
Description	Set the parity mode of the serial (via USB) interface.
Parameters	Name	Type	Range	Default
	<parity>	Discrete	EVEN|MARK|NONE|ODD|SPACE	NONE
Return	Query the parity mode of the serial interface.
Usage example	SYST:COMM:SER:PAR? / NONE
Related Commands	SYSTem:COMMunicate:ENABle

### SYSTem:CPU:INFOrmation:ETHernet:TYPE?
Syntax	SYSTem:CPU:INFOrmation:ETHernet:TYPE?
Description	This query returns the name of Ethernet controller installed on the control board (Arduino Shield). If simulator is used it returns “Simulator”.
Return	The information will be returned as a list of quoted strings.
Usage example	SYST:CPU:INFO:ETH:TYPE? / "W5500"

### SYSTem:CPU:INFOrmation:ONTime:LAST?
Syntax	SYSTem:CPU:INFOrmation:ONTime:LAST? [<channel>]
Description	This query returns time passed after last activation of the PSU. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CPU:INFO:ONT:LAST? / 15m

### SYSTem:CPU:INFOrmation:ONTime:TOTal?
Syntax	SYSTem:CPU:INFOrmation:ONTime:TOTal? [<channel>]
Description	This query returns total active time of the PSU. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Usage example	SYST:CPU:INFO:ONT:TOT? / "1h 45m"

### SYSTem:CPU:INFOrmation:TYPE?
Syntax	SYSTem:CPU:INFOrmation:TYPE?
Description	This query returns the name of CPU installed on the control board (Arduino Shield). If simulator is used it returns “Simulator”.
Usage example	SYST:CPU:TYPE? / "Due"
Related Commands	*IDN?

### SYSTem:CPU:MODel?
Syntax	SYSTem:CPU:MODel?
Description	This query returns the name of the control board (Arduino Shield). If simulator is used it returns “Simulator” and its version.
Usage example	If revision 1 Arduino Shield is detected: / SYST:CPU:MODel? / "Arduino, r5B12" / If firmware is running on simulator: / SYST:CPU:MODel? / "Simulator, M2.0"
Related Commands	SYSTem:CPU:OPTion?

### SYSTem:CPU:OPTion?
Syntax	SYSTem:CPU:OPTion?
Description	This query returns a list of all installed options on the control board. The information will be returned as a list of quoted strings.
Usage example	SYST:CPU:OPT? / "BPost, EEPROM, RTC, SDcard, Ethernet, Display, Watchdog"
Related Commands	SYSTem:CPU:MODel?

### SYSTem:DATE
Syntax	SYSTem:DATE {<year>,<month>,<day>} / SYSTem:DATE?
Description	Sets the date of the system clock (RTC). Specify the year, month, and day. / The self-test procedure compare date and time stored in RTC registers with values stored in the non-volatile memory (EEPROM). When the later is greater then former or any of them lost integrity (i.e. any of value is outside allowed range: for example seconds are higher then 60 or months are higher then 12, etc.) self-test will failed. The *TST? will return 1 and detailed report could be queried using the DIAGnostic:TEST? command. / The bit 3 (TIME) of the Questionable Status register will be set (see Section 3.4) if datetime self-test failed or datetime was never set.
Parameters	Name	Type	Range	Default
	<year>	Numeric (NR1)	2000 – 2099	–
	<month>	Numeric (NR1)	1 – 12	–
	<day>	Numeric (NR1)	1 – 31	–
Return	Query the current date of the system clock in YYYY, MM, DD format.
Usage example	SYST:DATE? / 2015, 10, 24
Related Commands	*TST? / DIAGnostic[:INFOrmation]:TEST? / SYSTem:TIME

### SYSTem:ERRor
Syntax	SYSTem:ERRor[:NEXT]?
Description	This query command reads and clear errors from the error queue. A record of up to 20 errors can be stored in the PSU’s error queue. See also “Error Messages” in Section 7. / Errors are retrieved in first-in-first-out (FIFO) order. The first error returned is the first error that was stored. The PSU beeps once each time an error is generated. The error queue is cleared when power has been off or after a *CLS command.
Return	SYSTem:ERRor[:NEXT]? queries and clears the error messages in the error queue. The query returns the number and content of the error message.
Usage example	SYST:ERR? / -113,"Undefined header"
Errors	If more than 20 errors have occurred, the last error stored in the queue (the most recent error) is replaced with: / -350,"Queue overflow" / No additional errors are stored until you remove errors from the queue.
Related Commands	*CLS / *RST / SYSTem:ERRor:COUNt

### SYSTem:ERRor:COUNt?
Syntax	SYSTem:ERRor:COUNt?
Description	This query command queries the error/event queue for the number of unread items. As errors and events may occur at any time, more items may be present in the queue at the time it is actually read.
Usage example	SYST:ERR:COUN? / 10
Related Commands	*CLS / *RST / SYSTem:ERRor[:NEXT]

### SYSTem:INHibit?
Syntax	SYSTem:INHibit?
Description	Use this command to query PSU’s inhibit state that is controlled with digital input (see [SOURce]:DIGital:PIN<n>:FUNCtion command). When PSU is in inhibited state, all channel’s output will be disabled (OFF) and triggered action (e.g. LIST) will be paused. / Channel output state command (OUTPut[:STATe]) can be used when PSU is in inhibited mode but it will not affect output (i.e. change it to enable state).
Return	Returns 0 if inhibit is not active or 1 if inhibit is active.
Usage example	SYST:INH? / 0
Related Commands	OUTPut[:STATe] {<bool>} [, <channel>] / [SOURce]:DIGital:PIN<n>:FUNCtion

### SYSTem:KLOCk
Syntax	SYSTem:KLOCk
Description	This command similar to the SYSTem:REMote command disables all front-panel options except for [Lock/Unlock] icon. You can push and hold the [Lock/Unlock] for a few seconds to unlock the front panel. The system password may be needed if it is set. / Lock state is saved in non-volatile memory. Therefore, the front panel remains locked even after AC power is cycled.
Usage example	SYST:KLOC
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:REMote

### SYSTem:LOCal
Syntax	SYSTem:LOCal
Description	This command places the PSU in the local mode during remote operation. All options on the front panel are fully functional.
Usage example	SYST:LOC
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:REMote / SYSTem:RWLock

### SYSTem:PASSword:CALibrate:RESet
Syntax	SYSTem:PASSword:CALibrate:RESet
Description	This command resets the calibration password to the firmware default setting, which is “eezpsu”. This command does not reset the system password.
Usage example	SYST:PASS:CAL:RES
Related Commands	CALibration[:MODE] / CALibration:PASSword:NEW

### SYSTem:PASSword:FPANel:RESet
Syntax	SYSTem:PASSword:FPANel:RESet
Description	This command resets the front panel lockout password to the firmware default setting, which is empty space (“”). This command does not reset the calibration password.
Usage example	SYST:PASS:FPAN:RES
Related Commands	SYSTem:KLOCk / SYSTem:PASSword:NEW

### SYSTem:PASSword:NEW
Syntax	SYSTem:PASSword:NEW {<old code>, <new>}
Description	Enter a new system password. To change the password, first unsecure the PSU using the old password. Then, the new code has to be entered. The calibration code may contain up to 16 characters over the remote interface. Minimum length is 4 characters. / The new password is automatically stored in non-volatile memory
Parameters	Name	Type	Range	Default
	<old>	Quoted string	0 to 16 characters	–
	<new>	Quoted string	4 to 16 characters	–
Usage example	SYST:PASS:NEW "","mypass2016"
Errors	122,"Invalid sys password" / 125,"Sys password too long" / 126,"Sys password too short"
Related Commands	SYSTem:KLOCk

### SYSTem:PON:OUTPut:DISable
Syntax	SYSTem:PON:OUTPut:DISable {<bool>} / SYSTem:PON:OUTPut:DISable?
Description	This command controls status off all channel outputs on power up. If enabled (ON), all outputs will be disabled regardless of what is stored in user profile selected for auto recall.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
Return	Query returns status of forced output disabling on power up.
Usage example	OUTP? / 1 / SYST:PON:OUTP 1 / (Restart) / OUTP? / 0
Related Commands	MEMory:STATe:RECall:AUTO / SYSTem:POWer

### SYSTem:POWer
Syntax	SYSTem:POWer {<bool>} / SYSTem:POWer?
Description	This command controls powering down and powering up sequence of the AC power inputs. The “Stand-by” indicator (LED_PWR, TLC5925 Out15) will be switched on when the PSU enters the Stand-by mode.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
Return	Query returns PSU power Stand-by status.
Usage example	SYST:POW ON / SYST:POW? / 1
Related Commands	*RST / *TST? / SYSTem:BEEP:STATe

### SYSTem:POWer:PROTection:TRIP
Syntax	SYSTem:POWer:PROTection:TRIP {<bool>} / SYSTem:POWer:PROTection:TRIP?
Description	Use this command to shut down (stand-by) the PSU when any of channel’s protection tripped. The SYSTem:POWer command has to be used to power the PSU on again.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
Return	Query returns status of shutdown when tripped mode.
Usage example	SYST:POW:PROT:TRIP? / 1
Related Commands	SYSTem:POWer

### SYSTem:REMote
Syntax	SYSTem:REMote
Description	Use this command to place the PSU into remote mode for serial (via USB) or Ethernet remote control. All front-panel options are disabled except for [Lock/Unlock] icon. You can push and hold the [Lock/Unlock] for a few seconds to unlock the front panel. The system password may be needed if it is set.
Usage example	SYST:REM
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:LOCal / SYSTem:RWLock

### SYSTem:RWLock
Syntax	SYSTem:RWLock
Description	Places the PSU in the remote mode for serial (via USB) or Ethernet remote control. This command is the same as SYSTem:REMote, except that all front panel options are disabled, including the [Lock/Unlock] icon.
Usage example	SYST:RWL
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:LOCal / SYSTem:REMote

### SYSTem:TEMPerature:PROTection[:HIGH][:LEVel]
Syntax	SYSTem:TEMPerature:PROTection[:HIGH][:LEVel] {<temperature>[, <sensor>]} / SYSTem:TEMPerature:PROTection[:HIGH][:LEVel]? [, {<sensor>}]
Description	Set the over-temperature protection (OTP) value in degrees Celsius (oC) of the selected temperature sensor. When the over-temperature protection function of the specified temperature sensor is enabled (SYSTem:TEMPerature:PROTection[:HIGH]:STATe), one of the following action will be performed when the temperature exceeds the over-temperature protection value currently set: / AUX – Switch off power of the main transformer and set bit 4 of the Questionable Status register / CH1, CH2 – Disable channel output (OUTPut OFF) and set bit 4 of the Questionable Instrument Isummary registerIf any of above mentioned temperature sensors cause over-temperature condition an error tone will also follow if beeper is enabled (see SYSTem:BEEPer:STATe). / SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped? command can be used to query whether over-temperature protection occurred on the selected temperature sensor.
Parameters	Name	Type	Range	Default
	<temperature>	Numeric (NR2)	0 – 100	70 for CH1, CH2, 50 for AUX
	<sensor>	Discrete	{[AUX]|CH1|CH2}	AUX
Return	Query the over-temperature protection (OTP) value of the selected temperature sensor.
Usage example	SYST:TEMP:PROT 50, AUX / SYST:TEMP:PROT? / 50
Related Commands	*RST / SYSTem:TEMPerature:PROTection[:HIGH]:STATe / SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped?

### SYSTem:TEMPerature:PROTection[:HIGH]:CLEar
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:CLEar [, {<sensor>}]
Description	This command clears the latched protection status when an over-temperature is detected. / All conditions that generate the fault must be removed before the latched status can be cleared. The output is restored to the state it was in before the fault condition occurred.
	Name	Type	Range	Default
	<sensor>	Discrete	{[AUX]CH1|CH2}	AUX
Usage example	SYST:TEMP:PROT:CLE
Related Commands	SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped

### SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME]
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME] {<delay>[, <sensor>]} / SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME]? [, {<sensor>}]
Description	This command sets the over-temperature protection delay. The over-temperature protection function will not be triggered during the delay time. After the delay time has expired, the over-temperature protection function will be active. / Programmed values can range from 0 to 300 seconds. See also Section 8.1
Parameters	Name	Type	Range	Default
	<delay>	Numeric (NR1)	0 – 300 seconds	10
	<sensor>	Discrete	{[AUX]CH1|CH2}	AUX
Return	The query returns programmed over-temperature protection delay.
Usage example	SYST:TEMP:PROT:DEL 30, CH2
Related Commands	*RST / SYSTem:TEMPerature:PROTection[:HIGH][:LEVel]

### SYSTem:TEMPerature:PROTection[:HIGH]:STATe
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:STATe {<bool>[, <sensor>]} / SYSTem:TEMPerature:PROTection[:HIGH]:STATe? [, {<sensor>}]
Description	This command enables or disables the over-temperature protection (OTP) function. The enabled state is ON (1); the disabled state is OFF (0). If the over-temperature protection function is enabled and the measured output power reach value set by [SOURce[<n>]]:POWer:PROTection[:LEVel] the output is disabled and the Questionable Condition status register OPP bit 10 is set.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
	<sensor>	Discrete	{[AUX]|CH1|CH2}	AUX
Return	The query command returns 0 if the current protection state is OFF, and 1 if the current protection state is ON.
Usage example	SYST:TEMP:PROT:STAT? CH1 / 0
Related Commands	*RST / SYSTem:TEMPerature:PROTection[:HIGH]:CLEar

### SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped?
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped? [, {<sensor>}]
Description	Query whether OTP occurred on the selected temperature sensor. When protection is tripped bit 4 (TEMPerature) of the Questionable Status register will be set (see Section 3.4). / The SYSTem:TEMPerature:PROTection[:HIGH]:CLEar command can be send to clear OTP condition caused by the selected temperature sensor.
Parameters	Name	Type	Range	Default
	<sensor>	Discrete	{[AUX]|CH1|CH2}	AUX
Return	This query returns a 1 if the protection circuit is tripped and a 0 if it is untripped.
Usage example	SYST:TEMP:PROT:TRIP? / 0
Related Commands	SYSTem:TEMPerature:PROTection[:HIGH]:CLEar / SYSTem:TEMPerature:PROTection[:HIGH]:STATe

### SYSTem:TIME
Syntax	SYSTem:TIME {<hours>,<minutes>,<seconds>} / SYSTem:TIME?
Description	Sets the time of the system clock (RTC). Specify the hours, minutes, and seconds. / The self-test procedure compare date and time stored in RTC registers with values stored in the non-volatile memory (EEPROM). When the later is greater then former or any of them lost integrity (i.e. any of value is outside allowed range: for example seconds are higher then 60 or months are higher then 12, etc.) self-test will failed. The *TST? will return 1 and detailed report could be queried using the DIAGnostic:TEST? command. / The bit 3 (TIME) of the Questionable Status register will be set (see Section 3.4) if datetime self-test failed or datetime was never set.
Parameters	Name	Type	Range	Default
	<hours>	Numeric (NR1)	0 – 23	–
	<minutes>	Numeric (NR1)	0 – 59	–
	<seconds>	Numeric (NR1)	0 – 59	–
Return	Query the current time of the system clock in HH, MM, SS format.
Usage example	SYST:TIME? / 15, 10, 33
Related Commands	*TST? / DIAGnostic[:INFOrmation]:TEST? / SYSTem:DATE

### SYSTem:TIME:DST
Syntax	SYSTem:TIME:DST {<rules>} / SYSTem:TIME:DST?
Description	Use this command to define daylight saving time (DST) rules used in your region. / Firmware v1.0 support limited number of region: Europe, US/Canada and Australia/New Zealand.
Parameters	Name	Type	Range	Default
	<rules>	Discrete	OFF|EU|USA|AUS	OFF
Return	Query the DTS rule name used for adjust time.
Usage example	SYST:TIME:DST EU
Related Commands	SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:NTP / SYSTem:TIME

### SYSTem:TIME:ZONE
Syntax	SYSTem:TIME:ZONE {<zone>} / SYSTem:TIME:ZONE?
Description	Use this command to define time zone as offset from GMT.
Parameters	Name	Type	Range	Default
	<zone>	Discrete	-12:00 to 14:00	–
Return	Query the time zone as offset from GMT in the following format: “[sign]hh:mm GMT”
Usage example	SYST:TIME:ZONE 1 / SYST:TIME:ZONE? / "+01:00 GMT"
Related Commands	SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:NTP / SYSTem:TIME

### SYSTem:VERSion?
Syntax	SYSTem:VERSion?
Description	This command returns the version of the SCPI (Standard Commands for Programmable Instruments) standard with which the instrument is in compliance
Return	The command returns a string in the form “YYYY.V”, where YYYY represents the year of the version and V represents a version for that year.
Usage example	SYST:VERS? / 1999.0