# Device-specific (unclassified) commands
The commands in this section are specific to the PSU, and so are not included in the 1999.0 version of the SCPI standard. However, these commands are designed with the SCPI standard in mind, and they follow all of the command syntax rules defined by the standard.
### APPLy
The APPLy command provides the most straightforward method to program the PSU remotely.
Syntax	APPLy {<channel>} [,<voltage>|MINimum|MAXimum|DEF[,<current>|MINimum|MAXimum|DEFault]] / APPLy? {<channel>}[, <param>]
Description	This command is a combination of the INSTrument:SELect (or INSTrument:NSELect), [SOURce[<n>]]:VOLTage and [SOURce[<n>]]:CURRent commands. / The APPLy changes the PSU’s output to the newly programmed values only if the programmed values are valid within the presently selected PSU operating range. An execution error will occur if the programmed values are not valid within the selected range. You can substitute MINimum, MAXimum, or DEFault in place of a specific value for the voltage and current parameters (see table below).
Parameters	Name	Type	Range	Default
	{<channel>}	Discrete	CH1|CH2	–
	<voltage>	Numeric	0 to MAXimum, MIN|DEF|MAX|UP|DOWN / The MAXimum value is dependent on the PSU voltage rating. See Section 8.1	–
	<current>	Numeric	0 to MAXimum, MIN|DEF|MAX|UP|DOWN / The MAXimum value is dependent on the PSU current rating. See Section 8.1	–
	<param>	Discrete	CURR|VOLT	–
Return	APPLy? query the voltage/current of the specified channel.
Usage example	Set the voltage and current of CH1 to 35.5V and 0.5A respectively: / APPL CH1,35.5,0.5 / Query the voltage and current settings of the first channel: / APPL? CH1 / CH1:50V/3A, 35.500, 0.500 / Query only current setting of the second channel: / APPL? CH2, CURR / 0.25
Errors	-221,"Power limit exceeded" / -222,"Data out of range"
Related Commands	INSTrument:NSELect / INSTrument[:SELect] / [SOURce[<n>]]:VOLTage[:LEVel][:IMMediate][:AMPLitude] / [SOURce[<n>]]:CURRent[:LEVel][:IMMediate][:AMPLitude] / TRIGger[:SEQuence]:SOURce

### DEBUg
Syntax	DEBUg?
Description	Use this command to collect various runtime information that can be used in debugging process.
Usage example	DEBU? / CH1 U_DAC = 0 / CH2 U_DAC = 32768 / CH1 U_MON = 0 / CH2 U_MON = 4095 / CH1 U_MON_DAC = 0 / CH2 U_MON_DAC = 16383 / CH1 I_DAC = 39321 / CH2 I_DAC = 6554 / CH1 I_MON = 0 / CH2 I_MON = 3276 / CH1 I_MON_DAC = 19660 / CH2 I_MON_DAC = 3276 / AUX TEMP = 373 / CH1 TEMP = 561 / CH2 TEMP = 373 / MAIN_LOOP_DURATION = 15998 16696 17409 / 0 0 0 / 1844 346805388 / ADC_COUNTER = 120 / 0 / 296
Related Commands	DIAGnostic[:INFOrmation]:ADC?