# Error messages
The system-defined error/event numbers are chosen on an enumerated (“1 of N”) basis. The SCPI-defined error/event numbers and the <error/event_description> portions of the full queue item are listed here. The first error/event described in each class (for example, -100, -200, -300, -400) is a “generic” error.
## Command Error [-199, -100]
An <error/event number> in the range [-199, -100] indicates that an syntax error has been detected by the PSU’s parser. The occurrence of any error in this class cause the command error bit (CME, bit 5) in the Standard Event Status Register (see Section 3.1) to be set.

Return string	Description
0,"No error"	The queue is completely empty. Every error/event in the queue has been read or the queue was purposely cleared by power-on, *CLS, etc.
−100,"Command error"	Generic syntax error.
-101,"Invalid character"	An invalid character was found in the command string. You may have inserted a character such as #, $, or % in the command keyword or within a parameter. / Example: / OUTP:STAT #ON
-103,"Invalid separator"	An invalid separator was found in the command string. You may have used a comma instead of a colon, semicolon, or blank space, or you may have used a blank space instead of a comma. / Example: / TRIG:SOUR,BUS
-104,"Data type error"	The wrong parameter type was found in the command string. You may have specified a number where a string was expected, or vice versa. / Example (password is not a quoted string): / CAL ON, 123
-108,"Parameter not allowed"	More parameters were received than expected for the command. You may have entered an extra parameter, or added a parameter to a command that does not accept a parameter. / Example: / INST CH1, CH2
-109,"Missing parameter"	Fewer parameters were received than expected for the command. You omitted one or more parameters that are required for this command. / Example: / APPL
-113,"Undefined header"	A command was received that is not valid for this PSU. You may have misspelled the command or it may not be a valid command. If you are using the short form of the command, remember that it may contain up to four letters. / Example: / MEASU:CURR?
-114,"Header suffix out of range"	The numeric suffix attached to a command header is not one of the allowable values. / Example: / STAT:QUES:INST:ISUM3?
-131,"Invalid suffix"	A suffix was incorrectly specified for a numeric parameter. You may have misspelled the suffix. / Example (use A instead of V): / VOLT 3A
-138,"Suffix not allowed"	A suffix was received following a numeric parameter which does not accept a suffix. / Example (SEC is not a valid suffix): / STAT:QUES:ENAB 10 SEC
-151,"Invalid string data"	An invalid character string was received. Check to see if you have enclosed the character string in single or double quotes. / Example: / DISP:TEXT ’ON

## Execution Error [-299, -200]
An <error/event number> in the range [-299, -200] indicates that an error has been detected by the PSU’s execution control block. The occurrence of any error in this class cause the execution error bit (EXE, bit 4) in the Standard Event Status Register (see Section 3.1) to be set. One of the following events has occurred:
- A <PROGRAM DATA> element following a header was evaluated by the device as outside of its legal input range or is otherwise inconsistent with the PSU’s capabilities.
- A valid program message could not be properly executed due to some PSU condition.
Execution errors will be reported by the PSU after rounding and expression evaluation operations have taken place. Rounding a numeric data element, for example, will not be reported as an execution error. Events that generate execution errors will not generate Command Errors, device-specific errors, or Query Errors; see the other error definitions in this section.

Return string	Description
-200,"Execution error"	This is the generic execution error when more specific error is not assigned in the case that command execution failed.
-222,"Data out of range"	A numeric parameter value is outside the valid range for the command. Example: / VOLT 166
-223,"Too much data"	A character string was received but could not be executed because the string length was more than 32 characters. This error can be generated by the CALibration:REMark command.
-224,"Illegal parameter value"	A discrete parameter was received which was not a valid choice for the command. You may have used an invalid parameter choice. / Example: / VOLT ON
-240,"Hardware error"	Command or query could not be executed because failure is detected during power-up self-test. Use *TST? command to query self-test results. See also Section 7.3.1
-241,"Option not installed"	Command or query could not be executed because of missing PSU hardware. / Example (remote sense cannot be activated when BP_OPTION is false): / OUTP:SENS ON
-242,"Channel fault detected"	POWERGOOD signal failure is detected on any of the installed channel. If such condition is happened the PSU will be immediately put into the stand-by mode.

## Device-Specific Error [-399, -300], [1, 32767]
An <error/event number> in the range [-399, -300] or [1, 32767] indicates that the PSU has detected an error which is not a command error, a query error, or an execution error; some PSU operations did not properly complete, possibly due to an abnormal hardware or firmware condition. These codes are also used for self-test response errors. The occurrence of any error in this class cause the device-specific error bit (DDE, bit 3) in the Standard Event Status Register (see Section 3.1) to be set.
Events that generate device-specific errors do not generate command errors, execution errors, or query errors; see the other error definitions in this section.

Return string	Description
-310,"System error"	Internal firmware error.
-330,"Self-test failed"	The PSU’s complete self-test failed from the remote interface (*TST? command). In addition to this error, more specific self-test errors are also reported. See also Section 7.3.1
-350,"Queue overflow"	The error queue is full because more than 16 errors have occurred. No additional errors are stored until you remove errors from the queue. The error queue is cleared when power has been turned off, or after a *CLS (clear status) command has been executed.
-363,"Input buffer overrun"	Input buffer overrun. Serial or Ethernet port input buffer overflows with data.
100,"Channel not found"	Non existing channel number is specified. / Example: / SOUR3:VOLT?
101,"Calibration mode is off"	Calibration is not enabled. The PSU will not accept calibration commands.
102,"Invalid cal password"	The calibration password is incorrect.
104,"Bad sequence of cal commands"	Calibration commands have not been entered in the proper sequence.
105,"Cal password too long"	A calibration password was received which contained more than 16 characters.
106,"Cal password too short"	A calibration password was received which contained less than 4 characters.
107,"Cal value out of range"	The specified calibration value (CALibration:CURRent:DATA or CALibration:VOLTage:DATA) is not valid for the present measurement function and range.
108,"Cal output disabled"	Calibration is aborted by sending the OUTP OFF command during an output calibration.
109,"Invalid cal data"	One or more calibration data have value that prohibit calibration parameters calculation. For example MID value is lower then MIN, or MIN is higher then MAX, or MID value is out of accepted tolerance that it cannot be predicted with newly calculated calibration parameters.
110,"Cal params missing or corrupted"	Calibration parameters activation initiated by CALibration:STATe ON, “<password>” failed because calibration has never been conducted or existing parameters are corrupted.
111,"No new cal data exists"	Attempt to save calibration data with the CAL:SAVE command while no new calibration data are entered.
150,"Power limit exceeded"	Product of voltage and current exceeds channel power limitation. For example if channel power limit is 160 W and the following sequence is executed: / VOLT 38 / CURR 4.4
201,"Cannot execute before clearing protection"	Command such as OUTP ON cannot be executed on the channel where one or more protections are tripped.
270,"CH1 ADC timeout detected"	Channel is switched off after 3 consecutive ADC timeouts are detected.

### Self-Test Error Messages
During power-up the PSU will start self-test sequence when communication with all SPI devices that is marked as installed will be established. Scope of self-test depends of device capability and it could vary from simple reading device registers and waiting for expecting response to more complex operation such as set DAC registers and read back set values using ADC (if ADC test passed). Every test failure will be announced by error beep, and one error message per failed test will be inserted into error queue. That also cause the device-specific error bit (DDE, bit 3) in the Standard Event Status Register (see Section 3.1) to be set.

Return string	Description
210,"CH1 IOEXP test failed"	Communication with Channel 1 I/O expander is not possible. I/O expander is not functional or SPI 10-pin ribbon cable is not connected.
211,"CH2 IOEXP test failed"	Communication with Channel 2 I/O expander is not possible. I/O expander is not functional or SPI 10-pin ribbon cable is not connected.
220,"CH1 ADC test failed"	Communication with Channel 1 ADC is not possible. Values written into various registers are not equal to the returned values or SPI 10-pin ribbon cable is not connected.
221,"CH2 ADC test failed"	Communication with Channel 2 ADC is not possible. Values written into various registers are not equal to the returned values or SPI 10-pin ribbon cable is not connected.
230,"CH1 DAC test failed"	Channel 1 DAC is not functional, communication is failed or difference between set and acquired DAC output value is too high when CPU sends a test voltage data to the DAC and converts the DAC output to digital data (using ADC I_SET input).
231,"CH2 DAC test failed"	Channel 2 DAC is not functional, communication is failed or difference between set and acquired DAC output value is too high when CPU sends a test voltage data to the DAC and converts the DAC output to digital data (using ADC I_SET input).
240,"External EEPROM test failed"	Non-volatile memory on the Arduino board checksum test failed.
250,"RTC test failed"	RTC on the Arduino board is not present or date or time values are not valid or last datetime stored into non-volatile memory is greater then RTC datetime.
260,"Ethernet test failed"	Ethernet controller on the Arduino board test failed.
630,"Fan test failed"

For example if SPI-bus cable is incidentally not connected or wrongly wired on channel 1 two errors will be generated and placed into error queue what can be check using the following command sequence:
SYST:ERR:COUN?
2
SYST:ERR?
210,"CH1 IOEXP test failed"
SYST:ERR?
220,"CH1 ADC test failed"