## SENSe
The SENSe commands are used .

SCPI command	Description
SENSe
:AHOur
:RESet	Resets the amp-hour (Ah) measurement to zero
:CURRent
[:DC]
RANGe[:UPPer] {<range>}	Selects a DC current measurement range
AUTO {<bool>}	Enables/disables seamless measurement auto ranging
:ELOG
:CURRent
[:DC]
:RANge[:UPPer] {<value>}	Selects the Elog current measurement range
:AUTO {<bool>}	Enables/disables Elog seamless measurement autoranging
:FUNCtion
:CURRent {<bool>}	Enables/disables current data logging
:VOLTage {<bool>}	Enables/disables voltage data logging
:PERiod {<time>}	Sets the integration time of an Elog measurement
:WHOur
:RESet	Resets the watt-hour (Wh) measurement to zero

### SENSe:AHOur:RESet
Not implemented yet
### SENSe:CURRent[:DC]:RANGe[:UPPer]
Syntax	SENSe:CURRent[:DC]:RANGe[:UPPer] {<range>} / SENSe:CURRent[:DC]:RANGe[:UPPer]?
Description	This command selects a DC current measurement range when Power board with multiple current ranges is installed (e.g. r5B12 that can be find out with the SYSTem:CHANnel:MODel? command). The entered value must be higher than the maximum current that you expect to measure. Units are in amperes. / If selected range is lower then measured current, when load is connected, an error -220 will be generated.
Return	The query command returns the maximum DC current that can be measured on the range that is presently set.
Parameters	Name	Type	Range	Default
	<range>	Numeric (NR2), discrete	0.5, 5, MIN|MAX	MAX
Usage example	SENS:CURR:RANG? / 0.5
Errors	-220,"Cannot set range"
Related Commands	SENSe:CURRent[:DC]:RANGe:AUTO / SYSTem:CHANnel:MODel?

### SENSe:CURRent[:DC]:RANGe:AUTO
Syntax	SENSe:CURRent[:DC]:RANGe:AUTO {<bool>} / SENSe:CURRent[:DC]:RANGe:AUTO?
Description	This command enables/disables current measurement autoranging. Autoranging provides seamless measurement ranging, which results in a wide dynamic range with no data being lost across ranges. / When enabled, this command overrides the SENSe:CURRent[:DC]:RANGe setting.
Return	The query command returns 0 if current measurement autoranging is OFF, and 1 if current measurement autoranging is ON.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	OFF
Usage example	SENS:CURR:RANG? / 0.5
Errors
Related Commands	SENSe:CURRent[:DC]:RANGe / SYSTem:CHANnel:MODel?

### SENSe:ELOG[:DC]:RANGe[:UPPer]
Not implemented yet
### SENSe:ELOG[:DC]:RANGe:AUTO
Not implemented yet
### SENSe:ELOG:FUNCtion:CURRent
Not implemented yet
### SENSe:ELOG:FUNCtion:VOLTage
Not implemented yet
### SENSe:ELOG:PERiod
Not implemented yet
### SENSe:WHOur:RESet
Not implemented yet