## OUTPut
The OUTPut subsystem controls the output, power-on, protection clear/reset and tracking state.

SCPI command	Description
OUTPut
[:STATe] {<bool>}	Controls the specified channel output state
:MODE?	Returns the channel mode of operation
:PROTection
:CLEar	Resets latched protection

### OUTPut[:STATe]
Syntax	OUTPut[:STATe] {<bool>} [, <channel>] / OUTPut[:STATe]? [<channel>]
Description	This command enables or disables the specified output channel(s). The enabled state is ON (1); the disabled state is OFF (0). The state of a disabled output is a condition of zero output voltage and zero source current. / Execution of this command also affects bit 10 (OE) and bit 11 (DP) of the Operation Instrument Isummary register (see Section 3.3.2). / Self-test operation initiated by *TST? command will put all PSU channels into disable state. / When the BP_OPTION is enabled and the channels are not grouped together (SYSTem:GROup commands scheduled for M3) this command sets LED indicators above binding posts in the following manner: / Turn on/off Out1 green indicator (LED_O1) when CH1 is selected / Turn on/off Out2 indicator (LED_O2) when CH2 is selected / Execution of the OUTP ON command on the channel with tripped one or more protection (OCP, OVP, OPP or OTP) will generate error 201. Use OUTPut:PROTection:CLEar command to clear all tripped protections. / This command also affects Sense+/Sense- indicators (LED_S1+, LED_S1-, LED_S2+, LED_S2-) controlled by the [SOURce:]VOLTage:SENSe[:SOURce] command.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	–
	<channel>	Discrete	CH1|CH2	–
Return	The query command returns 0 if the output is OFF, and 1 if the output is ON.
Usage example	OUTP ON, CH1 / OUTP? CH1 / 1
Errors	108,"Cal output disabled" / 201,"Cannot execute before clearing protection"
Related Commands	*TST / OUTPut:PROTection:CLEar / [SOURce:]VOLTage:SENSe[:SOURce]

### OUTPut:MODE?
Syntax	OUTPut:MODE? [, <channel>]
Description	This command simplify resolving a results that can be obtained reading the bit 8 (CV) and 9 (CC) of the read-only Instrument Isummary Operation Status register for a specific channel (see table in the Section 3.3.2). The PSU can works in one of the three output modes: / CV (Constant Voltage), when the output voltage equals the voltage setting value and the output current is determined by the load / CC (Constant Current), when the output current equals the current setting value and the output voltage is determined by the load and / UR (Unregulated) that is critical mode between CV and CC modes that could be noticed when the output voltage is close to the one end of the full scale (i.e. somewhere below 2V and above 48V for the 50V model that depends of the connected load). / The UR mode is not supported in software simulator.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Return	The query returns CV, CC or UR.
Usage example	Set output voltage to 20V and max. current, check that output voltage is as defined that indicate the constant voltage mode of operation: / VOLT 20; CURR MAX / MEAS:VOLT? / 20.0 / OUTP:MODE? / CV
Errors
Related Commands	STATus:OPERation:INSTrument:ISUMmary[<n>][:EVENt] / STATus:OPERation:INSTrument:ISUMmary[<n>]:CONDition / STATus:OPERation:INSTrument:ISUMmary[<n>]:ENABle

### OUTPut:PROTection:CLEar
Syntax	OUTPut:PROTection:CLEar [<channel>]
Description	This command clears the latched protection status that disables the output when an over-voltage, over-current or a power-limit condition is detected. / All conditions that generate the fault must be removed before the latched status can be cleared. The output is restored to the state it was in before the fault condition occurred.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2	–
Return	None
Usage example	The following command clears the latched protection status on all channels: / OUTP:PROT:CLE
Related Commands