Work in progress
# Error messages
The system-defined error/event numbers are chosen on an enumerated (“1 of N”) basis. The SCPI-defined error/event numbers and the <error/event_description> portions of the full queue item are listed here. The first error/event described in each class (for example, -100, -200, -300, -400) is a “generic” error.
## Command Error [-199, -100]
An <error/event number> in the range [-199, -100] indicates that an syntax error has been detected by the PSU’s parser. The occurrence of any error in this class cause the command error bit (CME, bit 5) in the Standard Event Status Register (see Section 3.1) to be set.

Return string	Description	M
0, "No error"	The queue is completely empty. Every error/event in the queue has been read or the queue was purposely cleared by power-on, *CLS, etc.	1
−100, "Command error"	Generic syntax error.	1
-101, "Invalid character"	An invalid character was found in the command string. You may have inserted a character such as #, $, or % in the command keyword or within a parameter. / Example: / OUTP:STAT #ON	1
-102, "Syntax error"	Invalid syntax was found in the command string. You may have inserted a blank space before or after a colon in the command header, or before a comma. / Example: / VOLT:LEV, 1	1
-103, Invalid separator	An invalid separator was found in the command string. You may have used a comma instead of a colon, semicolon, or blank space, or you may have used a blank space instead of a comma. / Example: / TRIG:SOUR,BUS	1
-104, "Data type error"	The wrong parameter type was found in the command string. You may have specified a number where a string was expected, or vice versa. / Example (password is not a quoted string): / CAL ON, 123	1
-105, "GET not allowed"	A Group Execute Trigger (GET) is not allowed within a command string.	3
-108, "Parameter not allowed"	More parameters were received than expected for the command. You may have entered an extra parameter, or added a parameter to a command that does not accept a parameter. / Example: / APPL? 10	1
-109, "Missing parameter"	Fewer parameters were received than expected for the command. You omitted one or more parameters that are required for this command. / Example: / APPL	1
-112, "Program mnemonic too long"	A command header was received which contained more than the maximum 12 characters allowed.	1
-113, "Undefined header"	A command was received that is not valid for this PSU. You may have misspelled the command or it may not be a valid command. If you are using the short form of the command, remember that it may contain up to four letters. / Example: / TRIGG:DEL 3	1
-114, "Header suffix out of range"	The numeric suffix attached to a command header is not one of the allowable values. / Example: / STAT:QUES:INST:ISUM3?	2
-115, "Unexpected number of parameters"	The number of parameters received does not correspond to the number of parameters expected. This is typically due an inconsistency with the number of instruments in the selected group (see section on INSTrument:DEFine:GROup).	?
-120, "Numeric data error"	An invalid number was specified for a numeric parameter. / Example: / APPL 1.0E+320000	2
-121, "Invalid character in number"	An invalid character was found in the number specified for a parameter value. / Example: / *ESE #B01010102	2
-123, "Exponent too large"	A numeric parameter was found whose exponent was larger than 32000.	2
-124, "Too many digits"	A numeric parameter was found whose mantissa contained more than 255 digits, excluding leading zeros.	2
-128, "Numeric data not allowed"	A numeric parameter was received but a character string was expected. / Example: / DISP:TEXT 123	2
-130, "Suffix error"	A suffix was incorrectly specified for a numeric parameter. You may have misspelled the suffix or the numeric parameter does not accept a suffix. / Example: / TRIG:DEL 0.5 SECS	2
-131, "Invalid suffix"	A suffix was incorrectly specified for a numeric parameter. You may have misspelled the suffix. / Example: / TRIG:DEL 0.5 SECS	1
-134, "Suffix too long"	A suffix for a numeric parameter contained too many characters.	2
-138, "Suffix not allowed"	A suffix was received following a numeric parameter which does not accept a suffix. / Example (SEC is not a valid suffix): / STAT:QUES:ENAB 10 SEC	1
-141, "Invalid character data"	Either the character data element contained an invalid character or the particular element received was not valid for the header.	2
-144, "Character data too long"	The character data element contained too many characters.	2
-148, "Character data not allowed"	A discrete parameter was received but a character string or a numeric parameter was expected. Check the list of parameters to verify that you have used a valid parameter type. / Example: / DISP:TEXT ON	2
-151, "Invalid string data"	An invalid character string was received. Check to see if you have enclosed the character string in single or double quotes. / Example: / DISP:TEXT ’ON	1
-158, "String data not allowed"	A character string was received but is not allowed for the command. Check the list of parameters to verify that you have used a valid parameter type. / Example: / TRIG:DEL ‘zero’	2
-160, "Block data error"	The PSU does not accept block data.	?
-170, "Expression error"	The PSU does not accept mathematical expressions.	1

## Execution Error [-299, -200]
An <error/event number> in the range [-299, -200] indicates that an error has been detected by the PSU’s execution control block. The occurrence of any error in this class cause the execution error bit (EXE, bit 4) in the Standard Event Status Register (see Section 3.1) to be set. One of the following events has occurred:
- A <PROGRAM DATA> element following a header was evaluated by the device as outside of its legal input range or is otherwise inconsistent with the PSU’s capabilities.
- A valid program message could not be properly executed due to some PSU condition.
Execution errors will be reported by the PSU after rounding and expression evaluation operations have taken place. Rounding a numeric data element, for example, will not be reported as an execution error. Events that generate execution errors will not generate Command Errors, device-specific errors, or Query Errors; see the other error definitions in this section.

Return string	Description	M
-200, "Execution error"	This is the generic execution error when more specific error is not assigned in the case that command execution failed.	1
-211, "Trigger ignored"	A *TRG was received but the trigger was ignored. The trigger source should be selected to the bus and the trigger subsystem should be initiated by the INIT[:IMM] command.	?
-213, "Init ignored"	An INITiate command was received but could not be executed because a measurement was already in progress. Send a device clear to halt a measurement in progress, and place the PSU in the “idle” state.	3
-221, "Settings conflict"	Indicates that a legal program data element was parsed but could not be executed due to the current device state.	?
-222, "Data out of range"	A numeric parameter value is outside the valid range for the command. Example: / VOLT 166	1
-223, "Too much data"	A character string was received but could not be executed because the string length was more than 32 characters. This error can be generated by the CALibration:STRing command.	1
-224, "Illegal parameter value"	A discrete parameter was received which was not a valid choice for the command. You may have used an invalid parameter choice. / Example: / VOLT ON	1
-240, "Hardware error"	Command or query could not be executed because failure is detected during power-up self-test. Use *TST? command to query self-test results. See also Section 7.3.1	1
-241, "Hardware missing"	Command or query could not be executed because of missing PSU hardware. / Example (remote sense cannot be activated when BP_OPTION is false): / OUTP:SENS ON	1
-242, "Channel fault detected"	Command or query could not be executed because POWERGOOD signal failure is detected	1
-252, "Missing media"	Command or query could not be executed because of a missing media (SDcard is not inserted).	3
-253, "Corrupt media"	Inserted SDcard is corrupted, not formatted or format is not recognized/supported.	3
-254, "Media full"	The media is full; there is no room on the inserted SDcard.	3
-256, "File name not found "	The file name on the device media was not found; for example, an attempt was made to read or copy a nonexistent file.	3
-257, "File name error"	The file name on the device media was in error; for example, an attempt was made to copy to a duplicate file name.	3
-258, "Media protected"	Indicates that a legal program command or query could not be executed because the media is protected (the lock slider is set in write-protect position).	3

## Device-Specific Error [-399, -300], [1, 32767]
An <error/event number> in the range [-399, -300] or [1, 32767] indicates that the PSU has detected an error which is not a command error, a query error, or an execution error; some PSU operations did not properly complete, possibly due to an abnormal hardware or firmware condition. These codes are also used for self-test response errors. The occurrence of any error in this class cause the device-specific error bit (DDE, bit 3) in the Standard Event Status Register (see Section 3.1) to be set.
Events that generate device-specific errors do not generate command errors, execution errors, or query errors; see the other error definitions in this section.

Return string	Description	M
-310		1
-330, "Self-test failed"	The PSU’s complete self-test failed from the remote interface (*TST? command). In addition to this error, more specific self-test errors are also reported. See also Section 7.3.1	1
-340, "Calibration failed"	General calibration error message that is generated when more specific information (Error codes 101 to 111) is not available.	3
-350, "Queue overflow"	The error queue is full because more than 16 errors have occurred. No additional errors are stored until you remove errors from the queue. The error queue is cleared when power has been turned off, or after a *CLS (clear status) command has been executed.	1
-363, “Input buffer overrun”		1
-365, "Time out error"	???	?
100, "Invalid channel number"	Non existing channel number is specified.	1
101, "Calibration mode is off"	Calibration is not enabled. The PSU will not accept calibration commands.	1
102, "Invalid cal password"	The calibration password is incorrect.	1
103, "Cal security disabled by jumper"	The calibration security feature has been disabled with a jumper inside the PSU (see Arduino shield board, JP2 position). When applicable, this error will occur at power-on to warn you that the PSU is unsecured.	3
104, "Bad sequence of cal commands"	Calibration commands have not been entered in the proper sequence.	1
105, "Cal password too long"	A calibration password was received which contained more than 16 characters.	1
106, "Cal password too short"	A calibration password was received which contained less than 4 characters.	1
107, "Cal value out of range"	The specified calibration value (CALibration:CURRent:DATA or CALibration:VOLTage:DATA) is not valid for the present measurement function and range.	1
108, "Cal output disabled"	Calibration is aborted by sending the OUTP OFF command during an output calibration.	1
109, "Invalid cal data"	One or more calibration data have value that prohibit calibration parameters calculation. For example MID value is lower then MIN, or MIN is higher then MAX, or MID value is out of accepted tolerance that it cannot be predicted with newly calculated calibration parameters.	1
110, "Cal params missing or corrupted"	Calibration parameters activation initiated by CALibration:STATe ON, “<password>” failed because calibration has never been conducted or existing parameters are corrupted.	1
111, "No new cal data exists"	Attempt to save calibration data with the CAL:SAVE command while no new calibration data are entered.	1
130, "Invalid sys password"	The system password is incorrect.	3
131, "Sys password too long"	A system password was received which contained more than 16 characters.	3
132, "Sys password too short"	A system password was received which contained less than 4 characters.	3
201, "Cannot execute before clearing protection"	Command such as OUTP ON cannot be executed on the channel where one or more protections are tripped.	1
302, "Option not installed"	The option that is programmed by this command is not installed.	3
800, "Outputs coupled by track system"	This error occurs when coupling the outputs if the track mode is already enabled. Tracking mode should be disabled first when coupling the outputs.	3
801, "Outputs coupled by trigger subsystem"	This error occurs when enabling the track mode if the outputs are already coupled. Outputs should be uncoupled first to enable the PSU to operate in the track mode.	3

### Self-Test Error Messages
During power-up the PSU will start self-test sequence when communication with all SPI devices that is marked as installed will be established. Scope of self-test depends of device capability and it could vary from simple reading device registers and waiting for expecting response to more complex operation such as set DAC registers and read back set values using ADC (if ADC test passed). Every test failure will be announced by error beep, and one error message per failed test will be inserted into error queue. That also cause the device-specific error bit (DDE, bit 3) in the Standard Event Status Register (see Section 3.1) to be set.

Return string	Description	M
200, "CH1 IOEXP test failed"	Add here description of the test	1
201, "CH2 IOEXP test failed"	Add here description of the test	1
210, "CH1 ADC test failed"	Add here description of the test	1
211, "CH2 ADC test failed"	Add here description of the test	1
220, "CH1 DAC test failed"	DAC on CH1 hardware is not functional. The MCU sends a reference voltage data to DAC and converts the DAC output to digital data (using ADC I_SET input) to see if the digital data is within a valid range.	1
230, "CH2 DAC test failed"	DAC on CH2 hardware is not functional. The MCU sends a reference voltage data to DAC and converts the DAC output to digital data (using ADC I_SET input) to see if the digital data is within a valid range.	1
240, "External EEPROM test failed"	Add here description of the test	1
250, "RTC test failed"	Add here description of the test	1
260, "Ethernet test failed"	Add here description of the test	1
270, "CH1 TEMP sensor failure"	Temperature sensor on CH1 is set as “installed” in system profile but no frequency is detected. Possible reason for that is disconnected NTC or V/F converter is not operational.	3
271, "CH2 TEMP sensor failure"	Temperature sensor on CH2 is set as “installed” in system profile but no frequency is detected. Possible reason for that is disconnected NTC or V/F converter is not operational.	3
280, "Display test failed"	MCU failed to establish communication with the TFT display controller.	3
281, "Input device test failed"	MCU failed to establish communication with the TFT touch screen controller.	3

## Query Error [-499, -400]
An <error/event number> in the range [-499, -400] indicates that the output queue control of the PSU has detected a problem with the message exchange protocol described in IEEE 488.2, chapter 6. The occurrence of any error in this class shall cause the query error bit (QYE, bit 2) in the Standard Event Status Register (see Section 3.1) to be set. One of the following is true:
- An attempt is being made to read data from the output queue when no output is either present or pending,
- Data in the output queue has been lost.
Events that generate query errors not generate command errors, execution errors, or device-specific errors; see the other error definitions in this section.

Return string	Description	M
-410, "Query INTERRUPTED"	A command was received which sends data to the output buffer, but the output buffer contained data from a previous command (the previous data is not overwritten). The output buffer is cleared when power has been turned off, or after a *RST command has been executed.	3
-420, "Query UNTERMINATED"	The PSU was addressed to talk (to send data over the interface) but a command has not been received which sends data to the output buffer. For example, you may have executed an APPLy command (which does not generate data) and then attempted an ENTER statement to read data from the remote interface.	3
-430, "Query DEADLOCKED"	A command was received which generates too much data to fit in the output buffer and the input buffer is also full. Command execution continues but all data is lost.	3
-440, "Query UNTERMINATED after indefinite response"	The *IDN? command must be the last query command within a command string. / Example: / *IDN?;:SYST:VERS?	3

## Power On Event [-599, -500]
An <error/event number> in the range [-599, -500] is used to report a power on event to the event/error queue. This event occurs when the PSU detects an off to on transition. This event also sets the power on bit, (PON, bit 7) of the Standard Event Status Register (see Section 3.1).

Return string	Description	M
-500, "Power on"	The PSU has detected an off to on transition. This function is accomplished by setting power on flag to true during initialization procedure. If PSU is regularly set to Stand by that flag will be reset and -501 message will be placed into the error queue.	3
-501, "Stand by"	The PSU will make available this message while entering Stand by mode.	3

## User Request Event [-699, -600]
An <error/event number> in the range [-699, -600] is used when the instrument wishes to report a user request event. This event occurs when the PSU detects the activation of a user request local control. This event also sets the user request bit (URQ, bit 6) of the Standard Event Status Register (see Section 3.1).

Return string	Description	M
-600, "User request"	The PSU has detected the activation of a user request local control. For example, a touch screen event is detected while PSU is in remote mode.	3

## Operation Complete Event [-899,-800]
An <error/event number> in the range [-899,-800] is used when the PSU wishes to report a 488.2 operation complete event to the error/event queue. This event occurs when PSU’s synchronization protocol, having been enabled by an *OPC command, completes all selected pending operations. This protocol is described in IEEE 488.2, section 12.5.2. This event also sets the operation complete bit (OPC, bit 0) of the Standard Event Status Register.

*OPC? does not set bit 0 nor does it enter any event in the error/event queue.

Return string	Description	M
-800, "Operation complete"	The instrument has completed all selected pending operations in accordance with the IEEE 488.2, 12.5.2 synchronization protocol.	3