# Device-specific (unclassified) commands
The commands in this section are device-specific to the PSU. They are not included in the 1999.0 version of the SCPI standard. However, these commands are designed with the SCPI standard in mind, and they follow all of the command syntax rules defined by the standard.
## APPLy
The APPLy command provides the most straightforward method to program the PSU over the remote interface.
Syntax	APPLy {<channel>} / [,<voltage>|MINimum|MAXimum|DEF[,<current>|MINimum|MAXimum|DEFault]] / APPLy? {<channel>}[, <param>]
Description	This command is a combination of INSTrument:SELect (or INSTrument:NSELect), [SOURce[<n>]]:VOLTage and [SOURce[<n>]]:CURRent commands. / The APPLy changes the PSU’s output to the newly programmed values only if the programmed values are valid within the presently selected range. An execution error will occur if the programmed values are not valid within the selected range. You can substitute MINimum, MAXimum, or DEFault in place of a specific value for the voltage and current parameters (see table below).
Parameters	Name	Type	Range	Default
	{<channel>}	Discrete	CH1|CH2	–
	<voltage>	Numeric	0 to maximum, MIN|DEF|MAX|UP|DOWN / The maximum value is dependent on the PSU voltage rating. See Section 8.1	–
	<current>	Numeric	0 to MAXimum, MIN|DEF|MAX|UP|DOWN / The maximum value is dependent on the PSU current rating. See Section 8.1	–
	<param>	Discrete	CURR|VOLT	–
Return	APPLy? query the voltage/current of the specified channel.
Usage example	Set the voltage and current of CH1 to 35.5V and 0.5A respectively: / APPL CH1,35.5,0.5 / Query the voltage and current settings of the first channel: / APPL? CH1 / CH1:50V/3A, 35.500, 0.500 / Query only current setting of the second channel: / APPL? CH2, CURR / 0.25
Errors	-221,"Power limit exceeded" / -222,"Data out of range"
Related Commands	INSTrument:NSELect / INSTrument[:SELect] / [SOURce[<n>]]:VOLTage[:LEVel][:IMMediate][:AMPLitude] / [SOURce[<n>]]:CURRent[:LEVel][:IMMediate][:AMPLitude]