## DISPlay
The DISPlay commands are used to set the display mode, turn on or off the front panel TFT display, select main page appearance, display and clear the text sent using a controller application. DISPlay is independent of, and does not modify, how data is returned to the controller application.

SCPI command	Description
DISPlay
:BRIGhtness {<value>}	Sets the intensity of the front panel TFT display
:VIEW {<mode>}	Sets the front panel TFT display main page appearance
[:WINdow]
[:STATe] {<bool>}	Sets the front panel TFT display state
:TEXT {<message>}	Displays a message on the front panel TFT display
:CLEar	Clear a message on the front panel TFT display

### DISPlay:BRIGhtness
Syntax	DISPlay:BRIGhtness {<brightness>} / DISPlay:BRIGhtness?
Description	Controls the intensity of the front panel TFT display. The range of the parameter is 1 to 20, where 20 is full intensity and 1 is fully blanked.
Parameters	Name	Type	Range	Default
	<value>	Numeric	1 – 20	20
Return	This query returns set front panel's TFT display brightness value.
Usage example	DISP:BRIG? / 20
Related Commands	*RST

### DISPlay:VIEW
Syntax	DISPlay:VIEW {<mode>} / DISPlay:VIEW?
Description	Use this command to set front panel TFT display main page appearance. The following modes are available: / 1 – Digits / 2 – Bar graph horizontal / 3 – Bar graph vertical / 4 – YT view
Parameters	Name	Type	Range	Default
	<mode>	Numeric	1 – 4|DEFault	DEFault
Return	This query returns set front panel's TFT display main page appearance numeric value.
Usage example	DISP:VIEW 2
Related Commands	*RST

### DISPlay[:WINdow][:STATe]
Syntax	DISPlay[:WINdow][:STATe] {<bool>} / DISPlay[:WINdow][:STATe]?
Description	Turn the front panel TFT display off or on. When the display is turned off, outputs are not sent to the display and all indicators are disabled except the Event view indicator. The display state is automatically turned on when you return to the local mode. / Press and hold the display for about a second to return to the LOCal from the REMote control.
Parameters	Name	Type	Range	Default
	<bool>	Discrete	ON|OFF|0|1	–
Return	DISPlay? query the front panel TFT display state. Returns 0 (OFF) or 1 (ON).
Usage example	DISP ON
Related Commands	SYSTem:LOCal / SYSTem:REMote

### DISPlay[:WINdow]:TEXT
Syntax	DISPlay[:WINdow]:TEXT {<text message>} / DISPlay[:WINdow]:TEXT?
Description	Display a message on the front panel TFT display. The PSU will display up to 32 characters in a message.
Parameters	Name	Type	Range	Default
	<text message>	Quoted string	–	–
Return	Query the message sent to the front panel TFT display and returns a quoted string.
Usage example	Send "Hello world" textual message: / DISP:TEXT "Hello world"

### DISPlay[:WINdow]:TEXT:CLEar
Syntax	DISPlay[:WINdow]:TEXT:CLEar
Description	Clear the message displayed on the front panel TFT display.
Return	None
Usage example	DISP:TEXT:CLE